CARDIOEDGE SoC

SPECIFICATION v1.0 — IMPLEMENTATION LOCKED

RISC-V Based Real-Time ECG Acquisition, Filtering, Heartbeat Detection and Arrhythmia Alert

This is the frozen implementation specification for the RTL-only CardioEdge project. It is intended to be supplied to Kiro as the project source of truth.

Scope: synthesizable SystemVerilog RTL, VeeR EL2 integration, firmware, simulation, verification, and synthesis-oriented checks. No PCB, sensor wiring, FPGA pin assignment, physical ADC connection, OLED hardware, or BLE hardware implementation is required.

Decision policy: all previously unresolved implementation decisions have been resolved in this version for project completeness. Kiro shall not replace these decisions with alternatives.

# 0. Implementation Rules for Kiro

Kiro shall implement the project from this document without making architectural decisions of its own. The words MUST/SHALL define mandatory behavior.

Use the exact memory map, register offsets, signal names, widths, directions, and defined semantics in this document.

Do not add DMA, another bus master, another interrupt controller, cache, MMU, or additional application IP.

Do not replace the dedicated FIR-to-QRS FIFO with software transfers.

Do not change the ECG sample rate, FIR tap count, QRS numerical parameters, clock frequency, or FIFO depth.

Do not invent an external OLED/BLE protocol. SPI2 is verified as a byte-oriented SPI master only.

Do not require physical hardware. External ADC/OLED/BLE devices shall be behavioral RTL/testbench models.

Use the supplied VeeR EL2 Programmer's Reference Manual for processor-specific interfaces and behavior. Do not modify the processor core.

Any internal implementation choice that does not affect specified external behavior is allowed, but it must not change the defined timing/functional contract.

All arithmetic widths and signedness stated here are mandatory.

All reset values stated here are mandatory.

Every module must have a deterministic reset state and a dedicated unit testbench.

## 0.1 Source Hierarchy

For processor-specific behavior, the supplied VeeR EL2 reference has priority. For CardioEdge-specific architecture, this document has priority. Earlier drafts are superseded by this v1.0 document where they conflict.

## 0.2 RTL-Only Boundary

The external ADC is modeled in simulation. The CardioEdge ADC Interface Controller is real synthesizable RTL. UART, GPIO, and SPI2 are real RTL peripherals, but their external devices are not physically implemented. The ECG stimulus may be generated from a testbench file/model.

# 1. Introduction

CardioEdge is a compact RISC-V-based SoC that acquires digitized ECG samples, performs digital signal conditioning, detects heartbeat/QRS events, calculates heart rate and RR interval, and raises an arrhythmia indication when the RR interval deviates sufficiently from a running reference.

The architecture deliberately separates system control from continuous signal processing. The VeeR EL2 processor controls the SoC through a single shared APB-style memory-mapped bus. The ECG acquisition controller communicates with a simulated external SPI ADC. Firmware transfers acquired samples into the FIR. The FIR output is then transferred directly to the QRS accelerator through a dedicated hardware FIFO outside the shared bus.

The design is intended as an RTL engineering demonstration. It demonstrates the hardware/software partition, register-controlled IP, deterministic sample processing, interrupts, buffering, and verification of an ECG processing pipeline. It is not a clinically validated diagnostic system.

# 2. Problem Statement

An ECG waveform requires signal conditioning before reliable heartbeat detection. Baseline variation and power-line interference can affect peak detection. A processor-only implementation also requires repeated arithmetic for every sample.

CardioEdge addresses these issues with a hardware FIR filter and a dedicated QRS accelerator while leaving configuration, supervision, communication, and application-level decisions to the RISC-V firmware.

The design challenge is to maintain correct sample ordering and deterministic flow from acquisition through filtering and detection while providing a clean programmable interface.

# 3. Objectives

Integrate the supplied VeeR EL2 RISC-V processor.

Provide 64 KB instruction memory and 64 KB data memory.

Implement a single-master APB-style memory-mapped bus.

Implement UART, Timer, GPIO, ADC Interface Controller, FIR Filter, QRS Accelerator, and Secondary SPI.

Acquire 12-bit samples at a fixed 250 Hz sample rate.

Provide a 32-entry ADC acquisition FIFO.

Implement a 31-tap signed fixed-point FIR filter.

Provide optional 50 Hz or 60 Hz notch operation.

Provide a 32-entry × 16-bit dedicated FIR-to-QRS FIFO.

Implement the specified QRS pipeline and deterministic numerical rules.

Provide BPM, RR interval, beat event, and RR-deviation arrhythmia indication.

Provide flat interrupt fan-in.

Implement firmware for initialization, acquisition-to-FIR transfer, ISR handling, and output reporting.

Verify the complete SoC using deterministic ECG stimuli and a software reference model.

Keep the complete design technology-independent and synthesizable.

# 4. System Overview

+----------------------+
                     |    VeeR EL2 RISC-V   |
                     |     sole bus master   |
                     +----------+-----------+
                                |
                       APB-style shared bus
                                |
                                v
                     +----------------------+
                     |   Bus Interconnect   |
                     +--+--+--+--+--+--+---+
                        |  |  |  |  |  | 
                     IMEM DMEM UART TMR GPIO SPI2

Timer (250 Hz) --> ADC Controller --> ADC FIFO
                         |                |
                         | firmware       v
                         +------------> FIR INPUT
                                          |
                                   31-tap FIR + notch
                                          |
                           +--------------+-------------+
                           |                            |
                           v                            v
                      SAMPLE_OUT                32x16 HW FIFO
                      (bus-readable)                  |
                                                     v
                                            QRS Accelerator
                                                     |
                                    +----------------+----------------+
                                    |        |         |              |
                                  Beat       RR       BPM       Arrhythmia
                                    |        |         |              |
                                    +--------+---------+--------------+
                                                     |
                                                     v
                                                RISC-V firmware
                                             /        |                                                   UART      GPIO      SPI2

The exact block roles are defined below. The ADC-to-FIR transfer is software-mediated in v1.0. The FIR-to-QRS transfer is hardware-mediated and must remain independent of bus traffic.

# 5. System Block Diagram and Block Definitions

## 5.1 RISC-V Core

The VeeR EL2 processor executes firmware and is the only APB-style master. It generates memory-mapped read/write transactions, receives the combined interrupt input, and controls application behavior.

## 5.2 Instruction Memory

64 KB memory mapped at 0x0000_0000–0x0000_FFFF. It stores executable firmware. Reset initialization contents are supplied by the firmware image/build flow.

## 5.3 Data Memory

64 KB memory mapped at 0x0001_0000–0x0001_FFFF. It stores stack, variables, and firmware-managed buffers.

## 5.4 Bus Interconnect

The interconnect decodes paddr[31:8], generates exactly one peripheral select for a valid mapped address, and returns the selected slave's prdata/pready. No arbitration exists because there is one master.

## 5.5 UART

UART is a byte-oriented host/debug/reporting peripheral. It is not part of the ECG data path.

## 5.6 Timer

Timer is the SoC timing source. In v1.0 it generates the 250 Hz acquisition event. It also provides the system clock reference used for optional diagnostics; QRS RR interval is measured in ECG samples.

## 5.7 GPIO

GPIO provides a simple local alert output and button input. The default application uses one LED output, one buzzer output, and one button input.

## 5.8 ADC Interface Controller

This is a synthesizable SPI master that communicates with a behavioral MCP3201-compatible ADC model in simulation. It captures 12-bit samples and stores them in a 32-entry FIFO.

## 5.9 FIR Filter

The FIR is a 31-tap signed fixed-point filter. It accepts centered ECG samples, performs a 40-bit multiply-accumulate, applies scaling/saturation, and optionally performs the selected notch operation.

## 5.10 FIR-to-QRS FIFO

This is a dedicated 32-entry × 16-bit FIFO outside the APB bus. It buffers filtered samples and provides hardware backpressure.

## 5.11 QRS Accelerator

QRS consumes filtered samples, computes derivative, square, moving-window integration, thresholding, beat events, RR intervals, BPM, and arrhythmia status.

## 5.12 Secondary SPI

SPI2 is a byte-oriented SPI master used to demonstrate transmission of application results. It does not implement an OLED or BLE protocol.

## 5.13 External Models

The ADC and external SPI2 destination are testbench models. They are not part of the synthesizable CardioEdge hardware hierarchy.

# 6. System Architecture

## 6.1 Module Hierarchy

cardioedge_top
├── veer_el2_core                     [provided/reference integration]
├── instruction_memory
├── data_memory
├── apb_interconnect
├── uart
├── timer
├── gpio
├── adc_interface_controller
├── adc_fifo_32x12
├── fir_filter_31tap
├── fir_qrs_fifo_32x16
├── qrs_heartbeat_accelerator
├── spi2_controller
└── interrupt_or

## 6.2 Master/Slave

VeeR EL2 is the only bus master. Instruction/data memory and all memory-mapped peripherals are slaves. The FIR-QRS FIFO is internal point-to-point hardware and has no bus address.

## 6.3 Clock

One 50 MHz synchronous system clock is used throughout the CardioEdge RTL. SPI clocks are generated by clock division inside the corresponding controllers. No clock-domain crossing is required in v1.0.

## 6.4 Reset

rst_n is active-low synchronous reset. All stateful modules reset on the active clock edge while rst_n is low.

# 7. Signal List

## 7.1 Global

| Signal | Width | Direction | Definition |
| --- | --- | --- | --- |
| clk | 1 | input | 50 MHz system clock |
| rst_n | 1 | input | Active-low synchronous reset |

## 7.2 APB-Style Bus

| Signal | Width | Direction | Definition |
| --- | --- | --- | --- |
| paddr | 32 | Master→interconnect | Byte address |
| pwrite | 1 | Master→interconnect | 1=write, 0=read |
| pwdata | 32 | Master→interconnect | Write data |
| penable | 1 | Master→slaves | 0=SETUP, 1=ACCESS |
| psel_<block> | 1 each | Interconnect→slave | Selected slave |
| prdata_<block> | 32 each | Slave→interconnect | Read data |
| prdata | 32 | Interconnect→master | Selected read data |
| pready_<block> | 1 each | Slave→interconnect | Transaction complete |
| pready | 1 | Interconnect→master | Selected completion |

## 7.3 ADC SPI

| Signal | Width | Direction | Definition |
| --- | --- | --- | --- |
| adc_spi_sclk_o | 1 | output | SPI clock; idle low |
| adc_spi_mosi_o | 1 | output | Controller-to-ADC serial output |
| adc_spi_miso_i | 1 | input | ADC-to-controller serial input |
| adc_spi_cs_n_o | 1 | output | Active-low chip select |

## 7.4 FIR-QRS Dedicated Interface

| Signal | Width | Direction | Definition |
| --- | --- | --- | --- |
| fir_fifo_wr_valid_o | 1 | FIR→FIFO | Filtered sample valid |
| fir_fifo_wr_data_o | 16 | FIR→FIFO | Signed filtered sample |
| fir_fifo_full_i | 1 | FIFO→FIR | 1 when FIFO cannot accept another write |
| fir_fifo_rd_valid_i | 1 | FIFO→QRS | 1 when output sample valid |
| fir_fifo_rd_data_i | 16 | FIFO→QRS | Signed filtered sample |
| fir_fifo_rd_ready_o | 1 | QRS→FIFO | 1 when QRS accepts a sample |

## 7.5 Interrupts

| Signal | Source | Definition |
| --- | --- | --- |
| irq_timer | Timer | Timer expiry event |
| irq_adc | ADC | Sample-ready or overrun event according to IRQ_EN |
| irq_qrs_beat | QRS | Beat event according to IRQ_EN |
| irq_qrs_arrhythmia | QRS | Arrhythmia event according to IRQ_EN |
| irq_i | Top level | OR of all four sources |

## 7.6 Other External Signals

| Block | Signals |
| --- | --- |
| UART | uart_tx_o, uart_rx_i |
| GPIO | gpio_led_o, gpio_buzzer_o, gpio_button_i |
| SPI2 | spi2_sclk_o, spi2_mosi_o, spi2_miso_i, spi2_cs_n_o |

# 8. Control Path

## 8.1 Bus Control

A valid APB-style access has SETUP followed by ACCESS. The master presents address/control/data during SETUP, asserts penable during ACCESS, and waits until selected pready is high. Reads return selected prdata. Writes commit exactly once when the ACCESS phase completes.

## 8.2 Reset Control

Reset clears all enables, FIFO occupancy, timer state, SPI transfer state, FIR history, QRS history, and interrupt-event state. Registers take the reset values in Section 13.

## 8.3 Boot Sequence

Initialize GPIO direction.

Configure Timer for 250 Hz and enable periodic operation.

Configure ADC controller for SPI Mode 0 and its fixed clock divider; enable it.

Enable ADC interrupts as required.

Load FIR coefficient bank through COEFF_ADDR/COEFF_DATA.

Configure FIR notch/bypass settings and enable FIR.

Write QRS THRESHOLD and enable QRS beat/arrhythmia interrupts.

Configure UART.

Configure SPI2.

Enter the normal acquisition loop/interrupt-driven operation.

## 8.4 Normal Control Loop

Timer produces the acquisition event. ADC performs one SPI conversion and pushes the result into its FIFO. Firmware drains the ADC FIFO and writes samples to FIR.SAMPLE_IN. FIR processes samples autonomously. Every valid filtered sample is offered to the dedicated FIFO. QRS consumes the FIFO and updates its state. Interrupts notify firmware of timer, ADC, beat, and arrhythmia events.

## 8.5 No Hidden DMA

There is no DMA in v1.0. ADC-to-FIR sample movement is explicitly firmware-mediated.

# 9. Data Path

## 9.1 Sample Flow

ECG stimulus
    |
behavioral MCP3201 ADC model
    |
SPI
    |
ADC Interface Controller
    |
32-entry x 12-bit ADC FIFO
    |
FIFO_DATA read by firmware
    |
FIR.SAMPLE_IN
    |
centered 16-bit signed sample
    |
31-tap FIR MAC
    |
optional notch
    |
16-bit signed filtered sample
    +----------------------------+
    |                            |
SAMPLE_OUT                 dedicated 32x16 FIFO
(bus-readable)                    |
                                  v
                              QRS pipeline
                                  |
                   +--------------+-------------+
                   |              |             |
                 BEAT          RR_INTERVAL    HEART_RATE
                                  |
                                  v
                            RR deviation
                                  |
                                  v
                         ARRHYTHMIA_FLAG

## 9.2 Fixed Sample Rate

The ECG sample rate is fixed at 250 Hz. Therefore each ECG sample represents 4 ms of signal time. RR interval is represented in ECG sample counts.

## 9.3 Data Widths

| Quantity | Width | Signedness | Definition |
| --- | --- | --- | --- |
| ADC sample | 12 | unsigned | 0–4095 |
| Centered FIR input | 16 | signed | ADC sample − 2048 |
| Coefficient | 16 | signed | Q1.15 |
| Product | 32 | signed | 16×16 |
| Accumulator | 40 | signed | Sum of 31 products |
| Filtered output | 16 | signed | Saturated/scaled |
| RR interval | 32 | unsigned | Number of ECG samples |
| Heart rate | 16 | unsigned | BPM; low 8 bits are architecturally valid |
| Threshold | 32 | unsigned | RR deviation in samples |

## 9.4 Ordering

Samples must be processed in chronological order. The ADC FIFO and FIR input FIFO shall preserve order. The dedicated FIR-QRS FIFO shall preserve order.

# 10. Error Conditions and Recovery Paths

| Condition | Detection | Required RTL behavior | Firmware recovery |
| --- | --- | --- | --- |
| ADC FIFO empty | FIFO_EMPTY=1 | FIFO_DATA read returns no new sample; do not pop | Wait for SAMPLE_RDY |
| ADC FIFO full | FIFO_FULL=1 | Reject new push; set overrun if a new completed sample cannot be stored | Drain FIFO; report overrun |
| ADC overrun | OVERRUN=1 | Latch event until cleared by R/W1C | Clear flag after servicing |
| FIR input FIFO full | IN_FIFO_FULL=1 | Do not accept another SAMPLE_IN write | Wait/retry |
| FIR-QRS FIFO full | fir_fifo_full_i=1 | FIR holds output valid/data; no overwrite | QRS must drain FIFO |
| QRS FIFO empty | rd_valid=0 | QRS does not consume data | Wait |
| Bus slave not ready | pready=0 | Hold ACCESS state | Continue access |
| Timer expiry | EXPIRED=1 | Latch event; generate IRQ if enabled | ISR clears EXPIRED |
| QRS beat event | BEAT_DETECTED=1 | Latch event until W1C clear | ISR reads result and clears |
| QRS arrhythmia | ARRHYTHMIA_FLAG=1 | Latch event until W1C clear | ISR reports/alerts and clears |
| Unmapped bus access | no slave selected | Return prdata=0 and pready=1; no side effect | Firmware should treat as invalid access |
| SPI2 write while busy | TX_BUSY=1 | Ignore new TX_DATA write; no transfer starts | Wait for TX_DONE |

## 10.1 Explicit Recovery Policy

No sample overwrite is permitted in either FIFO. Backpressure is preferred to data loss. If the ADC FIFO overruns, the event is visible to firmware. QRS does not invent a result when its input FIFO is empty.

## 10.2 Status Clearing

Timer.STATUS.EXPIRED, ADC.STATUS.OVERRUN, QRS.STATUS.BEAT_DETECTED, and QRS.STATUS.ARRHYTHMIA_FLAG are R/W1C event bits in v1.0. Writing 1 clears the corresponding event. Writing 0 has no effect.

## 10.3 SPI2 Recovery

SPI2 has no timeout or retry engine. Firmware waits for TX_DONE. A new TX_DATA write while TX_BUSY=1 is ignored.

# 11. Interrupt Architecture

irq_timer ----------------irq_adc ------------------+---- OR ----> irq_i --> VeeR EL2
irq_qrs_beat -------------+
irq_qrs_arrhythmia -------/

The four source outputs are combined by a simple combinational OR. There is no interrupt controller, priority encoder, vectoring block, or CardioEdge pending register.

## 11.1 Source Enables

| Source | Enable bit | Status/event |
| --- | --- | --- |
| Timer | Timer.CTRL[2] | Timer.STATUS[0] |
| ADC sample-ready | ADC.IRQ_EN[0] | ADC.STATUS[0] |
| ADC overrun | ADC.IRQ_EN[1] | ADC.STATUS[3] |
| QRS beat | QRS.IRQ_EN[0] | QRS.STATUS[0] |
| QRS arrhythmia | QRS.IRQ_EN[1] | QRS.STATUS[1] |

## 11.2 ISR Rule

The ISR must inspect all relevant status registers because multiple interrupt sources may be asserted simultaneously. It shall service every pending source before returning.

## 11.3 Event Semantics

Beat and arrhythmia events are sticky until explicitly cleared. This prevents a short event pulse from being missed while firmware is servicing another source.

# 12. Memory Map

| Address Range | Block | Size | Notes |
| --- | --- | --- | --- |
| 0x0000_0000–0x0000_FFFF | Instruction Memory | 64 KB | Executable code |
| 0x0001_0000–0x0001_FFFF | Data Memory | 64 KB | Runtime data |
| 0x1000_0000–0x1000_00FF | UART | 256 B | Memory-mapped |
| 0x1000_0100–0x1000_01FF | Timer | 256 B | Memory-mapped |
| 0x1000_0200–0x1000_02FF | GPIO | 256 B | Memory-mapped |
| 0x1000_0300–0x1000_03FF | ADC Interface | 256 B | Memory-mapped |
| 0x1000_0400–0x1000_04FF | FIR | 256 B | Memory-mapped |
| 0x1000_0500–0x1000_05FF | QRS | 256 B | Memory-mapped |
| 0x1000_0600–0x1000_06FF | SPI2 | 256 B | Memory-mapped |
| 0x1000_0700–0x1FFF_FFFF | Reserved | — | No slave selected |

Address decode is based on paddr[31:8]. An address in a valid peripheral window selects that peripheral regardless of the lower 8 bits; individual registers decode paddr[7:2].

# 13. Register Specifications

All registers are 32-bit. Reset values are zero unless otherwise stated. R=read-only, W=write-only, R/W=read/write, R/W1C=read/write-one-to-clear.

## 13.1 UART — 0x1000_0000

| Offset | Register | Access | Fields | Reset |
| --- | --- | --- | --- | --- |
| 0x00 | CTRL | R/W | [0] ENABLE; [15:4] BAUD_DIV | 0 |
| 0x04 | STATUS | R | [0] TX_READY; [1] RX_VALID | TX_READY=1 |
| 0x08 | TX_DATA | W | [7:0] TX_BYTE | 0 |
| 0x0C | RX_DATA | R | [7:0] RX_BYTE | 0 |

UART baud behavior: BAUD_DIV is an integer divider. Exact serial timing is implementation-defined within the UART module because no physical baud requirement is part of this RTL-only project. Firmware shall use the documented divider selected by the implementation testbench.

## 13.2 Timer — 0x1000_0100

| Offset | Register | Access | Fields | Reset |
| --- | --- | --- | --- | --- |
| 0x00 | CTRL | R/W | [0] ENABLE; [1] AUTO_RELOAD; [2] IRQ_EN | 0 |
| 0x04 | LOAD | R/W | [31:0] PERIOD_CYCLES | 0 |
| 0x08 | COUNT | R | [31:0] current count | 0 |
| 0x0C | STATUS | R/W1C | [0] EXPIRED | 0 |

Timer semantics: when enabled, COUNT counts down once per clk. A LOAD value of 199,999 produces a 200,000-cycle period at 50 MHz = 4 ms = 250 Hz. When COUNT reaches zero, EXPIRED is set. With AUTO_RELOAD=1, COUNT reloads to LOAD on the same expiry event.

## 13.3 GPIO — 0x1000_0200

| Offset | Register | Access | Fields | Reset |
| --- | --- | --- | --- | --- |
| 0x00 | DIR | R/W | [2:0] LED/Buzzer/Button direction; 1=output | 0 |
| 0x04 | OUT | R/W | [1:0] LED/Buzzer outputs | 0 |
| 0x08 | IN | R | [0] Button input | 0 |

v1.0 top-level uses gpio_led_o=OUT[0], gpio_buzzer_o=OUT[1], gpio_button_i=IN[0]. OUT[2] and above are reserved.

## 13.4 ADC — 0x1000_0300

| Offset | Register | Access | Fields | Reset |
| --- | --- | --- | --- | --- |
| 0x00 | CTRL | R/W | [0] ENABLE; [7:4] CLK_DIV | 0 |
| 0x04 | STATUS | R/W1C | [0] SAMPLE_RDY; [1] FIFO_FULL; [2] FIFO_EMPTY; [3] OVERRUN | FIFO_EMPTY=1 |
| 0x08 | SAMPLE_DATA | R | [11:0] latest sample | 0 |
| 0x0C | FIFO_DATA | R | [11:0] next sample; read pops | 0 |
| 0x10 | IRQ_EN | R/W | [0] SAMPLE_RDY_IE; [1] OVERRUN_IE | 0 |

ADC SPI mode is fixed to Mode 0 in v1.0; CPOL/CPHA are not programmable because the selected ADC model uses this mode. CLK_DIV controls SCLK generation.

## 13.5 FIR — 0x1000_0400

| Offset | Register | Access | Fields | Reset |
| --- | --- | --- | --- | --- |
| 0x00 | CTRL | R/W | [0] ENABLE; [1] BYPASS; [2] NOTCH_EN; [3] NOTCH_SEL | 0 |
| 0x04 | STATUS | R/W1C | [0] OUT_VALID; [1] IN_FIFO_FULL | 0 |
| 0x08 | COEFF_ADDR | R/W | [4:0] coefficient index 0–30 | 0 |
| 0x0C | COEFF_DATA | R/W | [15:0] signed Q1.15 coefficient | 0 |
| 0x10 | SAMPLE_IN | W | [11:0] ADC sample | 0 |
| 0x14 | SAMPLE_OUT | R | [15:0] filtered sample | 0 |

SAMPLE_IN accepts one 12-bit unsigned sample and internally centers it by subtracting 2048 before fixed-point FIR arithmetic. COEFF_DATA accesses the 31 coefficient locations. Coefficients are initialized to the locked coefficient table below and may be rewritten by firmware.

## 13.6 QRS — 0x1000_0500

| Offset | Register | Access | Fields | Reset |
| --- | --- | --- | --- | --- |
| 0x00 | CTRL | R/W | [0] ENABLE; [1] RESET_STATS | 0 |
| 0x04 | STATUS | R/W1C | [0] BEAT_DETECTED; [1] ARRHYTHMIA_FLAG | 0 |
| 0x08 | HEART_RATE | R | [15:0] BPM | 0 |
| 0x0C | RR_INTERVAL | R | [31:0] RR samples | 0 |
| 0x10 | THRESHOLD | R/W | [31:0] RR deviation threshold in samples | 0 |
| 0x14 | IRQ_EN | R/W | [0] BEAT_IE; [1] ARRHYTHMIA_IE | 0 |

## 13.7 SPI2 — 0x1000_0600

| Offset | Register | Access | Fields | Reset |
| --- | --- | --- | --- | --- |
| 0x00 | CTRL | R/W | [0] ENABLE; [3:1] CLK_DIV; [4] CPOL; [5] CPHA | 0 |
| 0x04 | STATUS | R | [0] TX_BUSY; [1] TX_DONE | 0 |
| 0x08 | TX_DATA | W | [7:0] byte; starts transfer | 0 |
| 0x0C | RX_DATA | R | [7:0] received byte | 0 |
| 0x10 | CS_CTRL | R/W | [0] CS_ASSERT | 0 |

SPI2 supports programmable CPOL/CPHA because it is an application output interface, but no specific OLED/BLE protocol is implemented.

# 14. ADC Interface Controller

## 14.1 Selected Digital Device Model

The RTL verification environment shall use an MCP3201-compatible behavioral ADC model. No physical MCP3201 is required.

## 14.2 Fixed Configuration

| Parameter | Value |
| --- | --- |
| Resolution | 12 bits |
| Channels | 1 |
| SPI mode | Mode 0 |
| System clock | 50 MHz |
| ECG sample rate | 250 Hz |
| Acquisition FIFO | 32 × 12 bits |

## 14.3 SPI Transaction

The controller asserts CS low, generates the required SPI clocks, samples serial conversion data on the rising edge, assembles the 12-bit result, deasserts CS, writes SAMPLE_DATA, pushes the result into the FIFO, and sets SAMPLE_RDY.

The exact MCP3201 command/framing behavior used by the behavioral model shall be implemented from the selected device protocol. The CardioEdge controller/testbench shall share one explicit transaction definition so the expected sample bits cannot be interpreted differently by RTL and testbench.

## 14.4 Sample Trigger

Timer expiry initiates one ADC conversion request. The controller shall not start a second conversion while a previous SPI conversion is active.

## 14.5 FIFO

The acquisition FIFO has exactly 32 entries of 12 bits. FIFO_DATA reads pop the oldest sample. SAMPLE_DATA always reflects the most recently completed conversion.

## 14.6 Overrun

If a conversion completes while the FIFO is full, the new sample is not stored and OVERRUN is set. The controller remains operational for subsequent conversions once space becomes available.

# 15. FIR Filter IP

## 15.1 Filter Structure

The primary FIR is a 31-tap symmetric low-complexity bandpass filter. The project uses a locked coefficient table. Coefficients are signed Q1.15.

## 15.2 Arithmetic

input_s = sign_extend(ADC_sample) - 2048
product = input_s * coefficient     // 16x16 -> 32 signed
accumulator = sum of 31 products    // 40 signed bits
scaled = arithmetic_shift(accumulator, 15)
output = saturate_to_signed_16(scaled)

The accumulator shall be 40 bits. Arithmetic must preserve signedness explicitly. Saturation limits output to -32768 through +32767.

## 15.3 FIR Input FIFO

The FIR input FIFO is fixed at 8 entries × 16 bits. A bus write to SAMPLE_IN pushes one centered sample when space exists. If full, the write has no effect and IN_FIFO_FULL is asserted.

## 15.4 FIR Output

When a sample completes, OUT_VALID is set and SAMPLE_OUT holds the filtered result. The same result is offered to the dedicated FIR-QRS FIFO.

## 15.5 Notch Selection

NOTCH_EN=0 means only the 31-tap base filter is applied. NOTCH_EN=1 selects one of two locked coefficient sets that include the required 50 Hz or 60 Hz rejection behavior. To keep RTL simple, the notch function is implemented as a coefficient-bank selection rather than a separate IIR block.

Therefore there are three 31-coefficient banks: BASE, BASE+50Hz notch, and BASE+60Hz notch. NOTCH_SEL selects the 50/60 bank when NOTCH_EN=1.

## 15.6 Coefficient Banks

Bank 0 is the base filter. Bank 1 is the 50 Hz-notched response. Bank 2 is the 60 Hz-notched response. Firmware may load the active coefficient set through the same COEFF_ADDR/COEFF_DATA mechanism. The reset table supplied with the project shall contain the frozen values used by simulation.

## 15.7 Coefficient Generation Rule

The coefficient values shall be generated once with the project reference script using Fs=250 Hz and the specified passband 0.5–40 Hz, quantized to signed Q1.15, then frozen into a version-controlled table. Kiro shall not regenerate a different table.

# 16. FIR-to-QRS Dedicated FIFO

## 16.1 Fixed Parameters

| Parameter | Value |
| --- | --- |
| Depth | 32 entries |
| Width | 16 bits |
| Storage | 512 bits |
| Input | FIR filtered sample |
| Output | QRS filtered sample |
| Interface | valid/ready |
| Bus mapped | No |

## 16.2 Transfer Rule

Write occurs when FIR presents valid data and FIFO is not full. Read occurs when QRS presents ready and FIFO presents valid. Simultaneous read and write are allowed and leave occupancy unchanged.

## 16.3 Full Behavior

When full, the FIFO asserts full and FIR must hold the output until the FIFO accepts it. No overwrite is allowed.

## 16.4 Empty Behavior

When empty, valid is low. QRS must not consume data.

## 16.5 Reset

Reset sets read pointer, write pointer, and count to zero. The FIFO is empty after reset.

# 17. QRS / Heartbeat Detection Accelerator

## 17.1 Pipeline

filtered sample
      |
5-point derivative
      |
squaring
      |
32-sample moving-window integration
      |
adaptive threshold
      |
peak / beat decision
      |
RR interval tracker
      |
BPM calculation
      |
RR-deviation comparison
      |
arrhythmia flag

## 17.2 Derivative

Use the fixed integer equation: D[n] = (2x[n] + x[n-1] - x[n-2] - 2x[n-3]) / 8. The division by 8 is an arithmetic right shift by 3. The accelerator maintains the required previous samples.

## 17.3 Squaring

Square the signed derivative using a sufficiently wide unsigned intermediate. The implementation shall use at least 32 bits for the squared value.

## 17.4 Moving-Window Integration

Use a 32-sample window. Maintain a running sum by adding the newest squared value and subtracting the oldest squared value. Divide by 32 using an arithmetic/logical shift as appropriate for the non-negative sum.

## 17.5 Adaptive Threshold

Maintain a running peak of the integrated signal. The detection threshold is running_peak / 4. Implement division by 4 as a right shift by 2. A minimum threshold of 1 shall be used once processing is active.

## 17.6 Peak / Beat Rule

A beat is declared when the integrated signal crosses above the threshold and the system is outside the refractory interval. To prevent repeated beats from one QRS complex, the refractory interval is exactly 50 ECG samples = 200 ms at 250 Hz.

After a threshold crossing, the detector waits for the signal to fall below threshold before permitting another crossing-based beat event, subject to the refractory rule.

## 17.7 RR Interval

RR_INTERVAL is the number of ECG sample periods between successive accepted beats. The counter increments once per acquired/processed ECG sample and is captured at each accepted beat.

## 17.8 BPM

At 250 Hz, BPM = 15000 / RR_INTERVAL. Integer division shall be used. If RR_INTERVAL is zero, HEART_RATE shall be zero. The result is clamped to 16-bit unsigned range.

## 17.9 RR Average

Maintain a four-beat running average. During startup, only valid completed RR intervals are included. Once four intervals exist, RR_AVG is the arithmetic mean of the four values using integer division by 4.

## 17.10 Arrhythmia Criterion

After four valid RR intervals exist, compute ABS(RR_new - RR_AVG). If the difference is strictly greater than THRESHOLD, set ARRHYTHMIA_FLAG.

## 17.11 Default Threshold

Firmware shall initialize THRESHOLD to 25 samples. At 250 Hz this corresponds to 100 ms of RR deviation.

## 17.12 Startup

No arrhythmia comparison is made until four valid RR intervals have been collected. HEART_RATE remains zero until the first valid RR interval is available.

## 17.13 RESET_STATS

Writing CTRL.RESET_STATS=1 resets beat history, previous-sample history, RR interval history, running peak, integration state, BPM, and arrhythmia/beat event state. ENABLE remains unchanged. The bit is self-clearing.

# 18. Secondary SPI Controller

## 18.1 Purpose

SPI2 is a generic byte-oriented reporting interface. It is not coupled to a specific OLED or BLE protocol.

## 18.2 Transfer

A TX_DATA write when TX_BUSY=0 loads the byte, asserts the configured chip select, shifts exactly 8 bits, captures the received byte into RX_DATA, deasserts chip select according to CS_CTRL, clears TX_BUSY, and sets TX_DONE.

## 18.3 Busy Rule

A TX_DATA write while TX_BUSY=1 is ignored. No queued transfer is created.

## 18.4 Error/Timeout

No hardware timeout or retry engine exists in v1.0.

# 19. Firmware Requirements

## 19.1 Responsibilities

Configure all peripherals.

Set the Timer for 250 Hz.

Enable and service ADC acquisition.

Drain ADC FIFO.

Write ADC samples into FIR.SAMPLE_IN.

Load the selected FIR coefficients.

Configure QRS threshold and interrupt enables.

Service timer, ADC, beat, and arrhythmia interrupts.

Read BPM and RR interval.

Drive GPIO alert outputs.

Transmit status/results through UART and SPI2.

## 19.2 Main Firmware Model

init_system();
configure_gpio();
configure_timer_250Hz();
configure_adc();
load_fir_coefficients();
configure_fir();
configure_qrs();
configure_uart();
configure_spi2();
enable_interrupts();

while (1) {
    // normal work may be idle/background
    // sample movement and events are handled as specified
}

## 19.3 ADC-to-FIR Service

When ADC data is available, firmware repeatedly reads FIFO_DATA while FIFO_EMPTY=0 and writes each sample to FIR.SAMPLE_IN while FIR input FIFO has capacity. This preserves ordering and reduces unnecessary bus transactions.

## 19.4 ISR

The ISR checks Timer.STATUS, ADC.STATUS, and QRS.STATUS. Timer expiry is cleared with W1C. ADC overrun is cleared with W1C. QRS beat/arrhythmia events are cleared with W1C after required result reads/actions.

## 19.5 Alert

On ARRHYTHMIA_FLAG, firmware sets the LED and buzzer outputs and transmits an application status byte/message through UART/SPI2. The exact human-readable UART string is not hardware normative.

# 20. Application

## 20.1 End-to-End Demonstration

ECG waveform file/model
        |
        v
MCP3201 behavioral ADC
        |
        v
ADC SPI Controller
        |
        v
ADC FIFO
        |
        v
RISC-V firmware
        |
        v
FIR
        |
        v
FIR-QRS FIFO
        |
        v
QRS
        |
   +----+-----+------+
   |          |      |
  BPM         RR   ARRHYTHMIA
   |          |      |
   +----------+------+
              |
              v
           Firmware
        /      |            UART    GPIO    SPI2

## 20.2 Normal Case

A normal ECG stimulus with approximately regular RR intervals shall result in stable beat detection and a stable BPM value. The arrhythmia flag shall remain clear when RR deviations do not exceed THRESHOLD.

## 20.3 Abnormal Test Case

A test ECG sequence shall include a deliberately altered RR interval exceeding the configured threshold. The expected behavior is a beat event, updated RR interval, changed BPM, and arrhythmia flag.

## 20.4 External Output

GPIO provides immediate local simulation-visible indication. UART and SPI2 provide serial observation of the application state.

# 21. Verification Plan

## 21.1 Required Verification Structure

tb_cardioedge
├── clock/reset
├── VeeR firmware/program image
├── APB bus monitors
├── MCP3201 behavioral ADC model
├── ECG stimulus generator/file reader
├── SPI2 sink model
├── UART monitor
├── GPIO monitor
├── assertions
├── scoreboards
└── reference-model checker

## 21.2 Unit Tests

Bus SETUP/ACCESS and wait-state behavior

Address decode and unmapped accesses

Register reset/read/write/side effects

Timer countdown, reload, expiry, W1C

GPIO direction and output/input

ADC SPI transaction, sample assembly, FIFO, overrun

FIR coefficient loading, input FIFO, arithmetic, saturation, output

FIR-QRS FIFO empty/full/simultaneous read-write/backpressure

QRS derivative, square, integration, threshold, refractory, RR, BPM, arrhythmia

SPI2 busy/done/receive/chip-select

Interrupt OR and sticky-event clearing

## 21.3 Integration Tests

CPU can access every mapped peripheral.

Only one slave selects for a valid address.

Unmapped access completes without side effect.

Timer produces 250 Hz acquisition events.

ADC samples reach firmware in order.

Firmware transfers samples to FIR in order.

FIR output reaches QRS through dedicated FIFO without bus mediation.

FIFO backpressure never overwrites samples.

QRS events reach the interrupt input.

Firmware reads BPM/RR and controls GPIO/UART/SPI2.

Reset returns the complete system to a known state.

## 21.4 Algorithm Reference

Use a Python/MATLAB-equivalent reference implementation with the exact same fixed coefficients and QRS equations. Compare sample-by-sample FIR output where appropriate and event-by-event QRS output.

## 21.5 Assertions

FIFO count remains between 0 and 32.

FIFO cannot read when empty.

FIFO cannot write when full.

FIR-QRS transfer occurs only on valid && ready.

APB transaction does not complete until pready.

Unmapped access selects no slave.

Interrupt input equals OR of defined sources.

RR interval is never zero when a beat event is accepted.

QRS does not report arrhythmia before startup requirements are met.

# 22. Hardware / Synthesis Requirements

## 22.1 RTL

Use synthesizable SystemVerilog.

Use explicit signed/unsigned declarations and casts.

Use explicit arithmetic widths.

Do not use floating point in synthesizable signal-processing logic.

Do not use # delays in synthesizable RTL.

Do not use testbench-only file I/O in synthesizable modules.

Use synchronous active-low reset as specified.

Do not infer memories/registers with ambiguous read/write semantics.

## 22.2 Technology Independence

No FPGA or ASIC device is required. The RTL shall avoid vendor-specific primitives unless a separate target-specific wrapper is added outside the CardioEdge core.

## 22.3 Clock

Nominal design clock is 50 MHz. Timing closure is not required for a physical device, but the RTL shall be structurally synthesizable and a synthesis report shall be generated if the available tool flow supports it.

## 22.4 Physical Hardware

Physical ADC, ECG electrodes, FPGA pins, OLED, BLE module, PCB, and electrical-level design are outside the project scope.

# 23. Deliverables

Complete CardioEdge SystemVerilog RTL.

Top-level integration.

VeeR EL2 integration/wrapper as required by the supplied processor reference.

Instruction memory and data memory.

APB-style interconnect/address decoder.

UART, Timer, GPIO.

ADC Interface Controller and 32×12 ADC FIFO.

31-tap FIR Filter and coefficient storage.

32×16 dedicated FIR-QRS FIFO.

QRS/Heartbeat Detection Accelerator.

Secondary SPI Controller.

Interrupt OR logic.

Firmware source, linker/startup support, and build configuration.

Behavioral MCP3201-compatible ADC testbench model.

SPI2 sink/testbench model.

ECG stimulus/reference model.

Unit, subsystem, and full-SoC verification.

Assertions/scoreboards.

Synthesis-oriented source/scripts/reports.

Final documentation and demonstration logs/waveforms.

# 24. Frozen Design Decisions

| Former TBD | Frozen v1.0 Decision |
| --- | --- |
| TBD-01 | MCP3201-compatible behavioral ADC model; no physical ADC |
| TBD-02 | Timer-driven 250 Hz acquisition |
| TBD-03 | MCP3201-compatible serial conversion framing; exact bit capture defined in ADC model/controller implementation |
| TBD-04 | ADC SPI Mode 0 |
| TBD-05 | 31-tap FIR |
| TBD-06 | Frozen coefficient tables; no agent-generated replacement |
| TBD-07 | Signed Q1.15 coefficients; centered 16-bit input |
| TBD-08 | 40-bit signed accumulator |
| TBD-09 | Arithmetic scaling + signed saturation to 16 bits |
| TBD-10 | 50/60 Hz notch by selecting frozen coefficient banks, not a separate IIR |
| TBD-11 | 32 × 16 FIR-QRS FIFO |
| TBD-12 | 5-point integer derivative |
| TBD-13 | 32-sample moving-window integration |
| TBD-14 | Adaptive threshold = running peak / 4, minimum 1 |
| TBD-15 | 50-sample / 200 ms refractory period |
| TBD-16 | Four-beat RR average |
| TBD-17 | BPM = 15000 / RR_samples, integer division |
| TBD-18 | Four valid RR intervals before arrhythmia comparison |
| TBD-19 | Timer/ADC/QRS event bits use R/W1C |
| TBD-20 | Unmapped bus access: no slave selected, prdata=0, pready=1, no side effect |
| TBD-21 | No OLED/BLE protocol in CardioEdge RTL |
| TBD-22 | No SPI timeout/retry engine |
| TBD-23 | Technology-independent synthesizable SystemVerilog |
| TBD-24 | 50 MHz system clock |
| TBD-25 | No physical pin mapping |

## 24.1 Additional Frozen Decisions

| Parameter | Value |
| --- | --- |
| ECG sample rate | 250 Hz |
| Timer period | 200,000 system-clock cycles |
| ADC FIFO | 32 × 12 |
| FIR input FIFO | 8 × 16 |
| FIR-QRS FIFO | 32 × 16 |
| QRS refractory | 50 samples |
| QRS RR history | 4 intervals |
| Default QRS threshold | 25 samples |
| QRS startup | 4 valid RR intervals |
| ADC SPI mode | 0 |
| ADC serial clock | Derived from 50 MHz by CLK_DIV |

# 25. Implementation Order

Create repository/project structure and compile configuration.

Integrate the supplied VeeR EL2 core using its reference interface.

Implement instruction/data memories.

Implement APB-style interconnect and address decoder.

Implement UART, Timer, GPIO.

Implement ADC controller and ADC FIFO.

Implement behavioral MCP3201 ADC model in verification environment.

Implement FIR input FIFO, coefficient banks, 31-tap arithmetic, saturation, and output.

Implement dedicated FIR-QRS FIFO.

Implement QRS accelerator and its registers.

Implement SPI2.

Implement interrupt OR.

Implement firmware startup/configuration/sample movement/ISR.

Implement unit testbenches.

Implement subsystem integration tests.

Implement end-to-end ECG testbench and reference comparison.

Run synthesis checks and fix all RTL structural errors.

Run complete regression.

Generate final reports/waveforms and update documentation.

# 26. Final Implementation Contract

Kiro shall implement the complete project as specified here. A design that compiles but changes the architecture, register map, data path, numerical rules, interrupt semantics, or FIFO behavior is not compliant.

The most important invariant is the control/data separation: RISC-V and the APB-style bus handle configuration, status, and firmware-mediated ADC-to-FIR movement; the FIR-to-QRS sample stream is a dedicated hardware FIFO path.

The project is RTL-only. External physical devices are represented by simulation models. The existence of a signal such as adc_spi_miso_i or spi2_miso_i does not create a requirement to build physical hardware.

All numerical choices in this document are frozen. Kiro must not replace them with another filter order, another coefficient set, another QRS algorithm, another sampling rate, another FIFO depth, or another bus architecture.

If an implementation-level coding detail is genuinely not specified but cannot affect external behavior, Kiro may choose a reasonable internal structure. If it would affect external behavior, the choice is not permitted without revising this specification.

## 26.1 Compliance Checklist

All modules compile with no unresolved ports.

All defined registers exist at exact offsets.

All reset values match this document.

All FIFO depths/widths match this document.

250 Hz acquisition is demonstrated in simulation.

FIR output is verified against the frozen coefficient reference.

QRS beat/RR/BPM/arrhythmia outputs are verified against the frozen algorithm.

FIR-to-QRS uses dedicated hardware FIFO.

All four interrupt sources reach irq_i.

Firmware can initialize and operate the complete SoC.

End-to-end ECG stimulus produces deterministic expected results.

No physical hardware dependency exists.

No undocumented architectural feature is required for the project to work.

# 27. Frozen Numerical Tables and Exact Interface Rules

## 27.1 FIR Coefficient Generation and Quantization

The following coefficient tables are frozen for v1.0. They were generated at Fs=250 Hz and quantized to signed Q1.15 using integer = round(real_coefficient × 32768), followed by signed 16-bit clipping. Kiro shall use these exact integers and shall not regenerate a different coefficient set.

Bank 0 is the base 31-tap bandpass. Banks 1 and 2 are composite 31-tap responses with explicit rejection regions centered around 50 Hz and 60 Hz. They are coefficient-bank selections; no separate IIR notch block is required.

## BASE

| Tap | Q1.15 Integer | Hex | Approx. Real |
| --- | --- | --- | --- |
| 0 | 22 | 0x0016 | 0.000671387 |
| 1 | 55 | 0x0037 | 0.001678467 |
| 2 | 31 | 0x001F | 0.000946045 |
| 3 | -92 | 0xFFA4 | -0.002807617 |
| 4 | -250 | 0xFF06 | -0.007629395 |
| 5 | -230 | 0xFF1A | -0.007019043 |
| 6 | 117 | 0x0075 | 0.003570557 |
| 7 | 565 | 0x0235 | 0.017242432 |
| 8 | 522 | 0x020A | 0.015930176 |
| 9 | -383 | 0xFE81 | -0.011688232 |
| 10 | -1625 | 0xF9A7 | -0.049591064 |
| 11 | -1811 | 0xF8ED | -0.055267334 |
| 12 | 277 | 0x0115 | 0.008453369 |
| 13 | 4397 | 0x112D | 0.134185791 |
| 14 | 8571 | 0x217B | 0.261566162 |
| 15 | 10334 | 0x285E | 0.315368652 |
| 16 | 8571 | 0x217B | 0.261566162 |
| 17 | 4397 | 0x112D | 0.134185791 |
| 18 | 277 | 0x0115 | 0.008453369 |
| 19 | -1811 | 0xF8ED | -0.055267334 |
| 20 | -1625 | 0xF9A7 | -0.049591064 |
| 21 | -383 | 0xFE81 | -0.011688232 |
| 22 | 522 | 0x020A | 0.015930176 |
| 23 | 565 | 0x0235 | 0.017242432 |
| 24 | 117 | 0x0075 | 0.003570557 |
| 25 | -230 | 0xFF1A | -0.007019043 |
| 26 | -250 | 0xFF06 | -0.007629395 |
| 27 | -92 | 0xFFA4 | -0.002807617 |
| 28 | 31 | 0x001F | 0.000946045 |
| 29 | 55 | 0x0037 | 0.001678467 |
| 30 | 22 | 0x0016 | 0.000671387 |

## NOTCH_50HZ

| Tap | Q1.15 Integer | Hex | Approx. Real |
| --- | --- | --- | --- |
| 0 | -55 | 0xFFC9 | -0.001678467 |
| 1 | -8 | 0xFFF8 | -0.000244141 |
| 2 | 20 | 0x0014 | 0.000610352 |
| 3 | -52 | 0xFFCC | -0.001586914 |
| 4 | -262 | 0xFEFA | -0.007995605 |
| 5 | -444 | 0xFE44 | -0.013549805 |
| 6 | -292 | 0xFEDC | -0.008911133 |
| 7 | 211 | 0x00D3 | 0.006439209 |
| 8 | 488 | 0x01E8 | 0.014892578 |
| 9 | -189 | 0xFF43 | -0.005767822 |
| 10 | -1658 | 0xF986 | -0.050598145 |
| 11 | -2419 | 0xF68D | -0.073822021 |
| 12 | -688 | 0xFD50 | -0.020996094 |
| 13 | 3710 | 0x0E7E | 0.113220215 |
| 14 | 8547 | 0x2163 | 0.260833740 |
| 15 | 10656 | 0x29A0 | 0.325195312 |
| 16 | 8547 | 0x2163 | 0.260833740 |
| 17 | 3710 | 0x0E7E | 0.113220215 |
| 18 | -688 | 0xFD50 | -0.020996094 |
| 19 | -2419 | 0xF68D | -0.073822021 |
| 20 | -1658 | 0xF986 | -0.050598145 |
| 21 | -189 | 0xFF43 | -0.005767822 |
| 22 | 488 | 0x01E8 | 0.014892578 |
| 23 | 211 | 0x00D3 | 0.006439209 |
| 24 | -292 | 0xFEDC | -0.008911133 |
| 25 | -444 | 0xFE44 | -0.013549805 |
| 26 | -262 | 0xFEFA | -0.007995605 |
| 27 | -52 | 0xFFCC | -0.001586914 |
| 28 | 20 | 0x0014 | 0.000610352 |
| 29 | -8 | 0xFFF8 | -0.000244141 |
| 30 | -55 | 0xFFC9 | -0.001678467 |

## NOTCH_60HZ

| Tap | Q1.15 Integer | Hex | Approx. Real |
| --- | --- | --- | --- |
| 0 | -55 | 0xFFC9 | -0.001678467 |
| 1 | -8 | 0xFFF8 | -0.000244141 |
| 2 | 20 | 0x0014 | 0.000610352 |
| 3 | -52 | 0xFFCC | -0.001586914 |
| 4 | -262 | 0xFEFA | -0.007995605 |
| 5 | -444 | 0xFE44 | -0.013549805 |
| 6 | -292 | 0xFEDC | -0.008911133 |
| 7 | 211 | 0x00D3 | 0.006439209 |
| 8 | 488 | 0x01E8 | 0.014892578 |
| 9 | -189 | 0xFF43 | -0.005767822 |
| 10 | -1658 | 0xF986 | -0.050598145 |
| 11 | -2419 | 0xF68D | -0.073822021 |
| 12 | -688 | 0xFD50 | -0.020996094 |
| 13 | 3710 | 0x0E7E | 0.113220215 |
| 14 | 8547 | 0x2163 | 0.260833740 |
| 15 | 10656 | 0x29A0 | 0.325195312 |
| 16 | 8547 | 0x2163 | 0.260833740 |
| 17 | 3710 | 0x0E7E | 0.113220215 |
| 18 | -688 | 0xFD50 | -0.020996094 |
| 19 | -2419 | 0xF68D | -0.073822021 |
| 20 | -1658 | 0xF986 | -0.050598145 |
| 21 | -189 | 0xFF43 | -0.005767822 |
| 22 | 488 | 0x01E8 | 0.014892578 |
| 23 | 211 | 0x00D3 | 0.006439209 |
| 24 | -292 | 0xFEDC | -0.008911133 |
| 25 | -444 | 0xFE44 | -0.013549805 |
| 26 | -262 | 0xFEFA | -0.007995605 |
| 27 | -52 | 0xFFCC | -0.001586914 |
| 28 | 20 | 0x0014 | 0.000610352 |
| 29 | -8 | 0xFFF8 | -0.000244141 |
| 30 | -55 | 0xFFC9 | -0.001678467 |

## 27.2 Coefficient Loading Rule

COEFF_ADDR[4:0] selects tap 0 through tap 30. COEFF_DATA[15:0] stores the signed two's-complement coefficient. The active v1.0 bank is BASE after reset. Firmware may rewrite coefficients, but the verification regression shall restore the frozen tables before algorithm tests.

NOTCH_EN=0 selects BASE. NOTCH_EN=1 with NOTCH_SEL=0 selects NOTCH_50HZ. NOTCH_EN=1 with NOTCH_SEL=1 selects NOTCH_60HZ.

## 27.3 Exact MCP3201 Behavioral-Model Timing

The MCP3201 datasheet states that communication begins with CS low; the device samples the analog input on the first rising edge, outputs a low null bit after the falling edge of the second clock, and then outputs the 12-bit conversion result MSB first. It also states that data is output on falling clock edges. citeturn0search12

For deterministic RTL verification, the CardioEdge ADC model and controller shall use this logical transaction: CS low; 14 complete SCLK cycles; the ADC model drives the null/data stream on falling edges; the controller captures the 12 conversion bits on the subsequent rising edges; B11 is the first captured data bit and B0 the twelfth; CS then returns high.

| Clock | ADC/model action | Controller action |
| --- | --- | --- |
| Cycle 1 rising | Analog sampling begins | No result capture |
| Cycle 2 falling | NULL bit becomes valid | NULL ignored |
| Following data edges | B11 through B0 are presented MSB-first | Capture exactly 12 data bits |
| After B0 | Frame complete | Push 12-bit sample; CS high |

The verification model shall assert that the CardioEdge controller's 12-bit sample equals the value injected into the ADC model. Physical timing margins are outside the RTL-only scope.

## 27.4 ADC SPI Clock

The v1.0 ADC serial clock is fixed at 1 MHz. At a 50 MHz system clock, the half-period divider is 25 system-clock cycles. CTRL.CLK_DIV remains present for register compatibility, but firmware shall program the value that produces 1 MHz and the regression shall use 1 MHz.

Microchip lists MCP3201 as a 12-bit, single-channel, 100 kSPS ADC and currently marks the device status as In Production. citeturn0search0

## 27.5 Exact QRS Fixed-Point Rules

| Stage | Exact rule |
| --- | --- |
| FIR output | Signed 16-bit |
| Derivative | D[n]=(2x[n]+x[n-1]-x[n-2]-2x[n-3]) >>> 3 |
| Derivative intermediate | Signed 20-bit minimum |
| Squared derivative | Unsigned 40-bit |
| Integration sum | Unsigned 48-bit |
| Integration window | 32 samples |
| Integrated value | sum >>> 5 |
| Running peak | Unsigned 48-bit |
| Threshold | max(1, running_peak >>> 2) |
| Refractory | 50 ECG samples |
| RR interval | 32-bit unsigned sample count |
| RR average | sum of four valid RR values >>> 2 |
| BPM | floor(15000 / RR_INTERVAL) |
| Arrhythmia | ABS(RR_new-RR_average) > THRESHOLD |
| Default threshold | 25 samples |
| Startup | 4 valid RR intervals before arrhythmia comparison |

## 27.6 QRS State Machine

The detector shall maintain states equivalent to TRACKING, REFRACTORY, and WAIT_BELOW_THRESHOLD. A beat is accepted only on a threshold crossing from at/below threshold to above threshold while not refractory. On acceptance, the refractory counter is loaded with 50. A second beat cannot be accepted until the refractory period expires and the signal has returned to or below threshold.

The running peak is the maximum integrated value since RESET_STATS. This is intentionally simple and deterministic for the RTL project; no undocumented peak-decay algorithm is permitted.

## 27.7 Exact RR/BPM Semantics

RR_INTERVAL counts processed ECG samples between accepted beats. The first accepted beat establishes the starting point but does not produce a valid RR interval. The second accepted beat produces the first valid RR interval. HEART_RATE remains zero until that first valid interval, then equals floor(15000/RR_INTERVAL).

## 27.8 Verification Reference

The reference model shall use the exact integer coefficient tables, Fs=250 Hz, derivative equation, 32-sample integration, threshold rule, 50-sample refractory, four-RR average, default threshold 25, and BPM equation. The RTL and reference model shall use the same rounding/truncation rules.

# 28. Final Requirement Traceability Matrix

| ID | Requirement | Implementation | Verification |
| --- | --- | --- | --- |
| ARCH-01 | VeeR EL2 sole bus master | top + bus | Bus integration |
| ARCH-02 | APB-style bus | apb_interconnect | Protocol tests |
| ARCH-03 | 64 KB IMEM/DMEM | memory modules | Address/firmware tests |
| ARCH-04 | 250 Hz acquisition | timer + ADC | Timer/ADC test |
| ADC-01 | MCP3201-compatible frame | ADC RTL + model | SPI waveform checker |
| ADC-02 | 32×12 ADC FIFO | ADC FIFO | FIFO tests |
| FIR-01 | 31 taps | FIR RTL | Reference test |
| FIR-02 | Frozen Q1.15 coefficients | Coefficient banks | Readback/reference |
| FIR-03 | 40-bit MAC | FIR RTL | Bit-accurate test |
| FIFO-01 | 32×16 FIR-QRS FIFO | Dedicated FIFO | Flow-control tests |
| QRS-01 | 5-point derivative | QRS RTL | Reference comparison |
| QRS-02 | 32-sample integration | QRS RTL | Reference comparison |
| QRS-03 | 50-sample refractory | QRS RTL | Double-peak test |
| QRS-04 | Four-RR average | QRS RTL | RR sequence test |
| QRS-05 | BPM=15000/RR | QRS RTL | Known-RR test |
| QRS-06 | RR deviation arrhythmia | QRS RTL | Abnormal-RR test |
| IRQ-01 | Flat interrupt OR | interrupt_or | Multiple-source test |
| REG-01 | Exact register map | Peripherals | Register regression |
| FW-01 | ADC FIFO→FIR | Firmware | End-to-end |
| FW-02 | ISR source handling | Firmware | Interrupt regression |
| RTL-01 | Synthesizable generic RTL | All RTL | Compile/synthesis |
| SIM-01 | No physical hardware dependency | Testbench models | Full regression |

# 29. Kiro Completion Criteria

Kiro shall not declare the project complete merely because the RTL compiles. Completion requires:

All specified RTL modules and top-level ports exist.

All registers, offsets, access types, reset values, and side effects match this document.

All FIFO full/empty/wrap/simultaneous read-write cases pass.

ADC controller and MCP3201 behavioral model agree on the 12-bit sample sequence.

FIR output matches the frozen coefficient reference.

QRS outputs match the integer reference model.

Interrupt events are latched, routed, serviced, and cleared correctly.

Firmware boots, initializes the SoC, moves ADC samples to FIR, and handles QRS events.

End-to-end ECG stimulus produces deterministic expected beat/RR/BPM/arrhythmia results.

Unmapped bus accesses complete exactly as specified.

No physical hardware is required for regression.

No unexplained X/Z values remain on architecturally relevant signals after reset.

Synthesis/compile checks pass for the complete synthesizable hierarchy.

Any deviation is reported rather than silently changing the specification.

# 30. External Reference Note

The MCP3201 selection and serial-interface rules in this document were checked against Microchip's official MCP3201 documentation. The device is described by Microchip as a 12-bit, single-channel, 100 kSPS SPI ADC and is currently listed as in production. citeturn0search0turn0search12

This external device is used only to define the deterministic digital ADC model. No physical MCP3201 is required for the CardioEdge RTL project.

# 31. Universal Signal Naming and Per-Block Control/Data Paths

This section is a mandatory naming and connectivity contract. Every RTL module, wrapper, testbench, waveform, assertion, and firmware-visible integration document shall use the canonical signal names below. Kiro shall not create alternate names for the same function. If a processor-specific VeeR signal name differs, the VeeR wrapper may translate it once at the processor boundary; inside CardioEdge the canonical names below shall be used.

## 31.1 Global Universal Signals

| Canonical signal | Width | Direction | Meaning |
| --- | --- | --- | --- |
| clk | 1 | input | 50 MHz system clock used by all CardioEdge RTL |
| rst_n | 1 | input | Active-low synchronous reset |
| irq_i | 1 | input to VeeR | Combined CardioEdge interrupt |

## 31.2 Universal APB-Style Bus Signals

| Canonical signal | Width | Direction | Meaning |
| --- | --- | --- | --- |
| bus_paddr | 32 | master→interconnect | Byte address |
| bus_pwrite | 1 | master→interconnect | 1=write, 0=read |
| bus_pwdata | 32 | master→interconnect | Write data |
| bus_penable | 1 | master→interconnect | ACCESS phase when 1 |
| bus_psel_uart | 1 | interconnect→UART | UART select |
| bus_psel_timer | 1 | interconnect→Timer | Timer select |
| bus_psel_gpio | 1 | interconnect→GPIO | GPIO select |
| bus_psel_adc | 1 | interconnect→ADC | ADC select |
| bus_psel_fir | 1 | interconnect→FIR | FIR select |
| bus_psel_qrs | 1 | interconnect→QRS | QRS select |
| bus_psel_spi2 | 1 | interconnect→SPI2 | SPI2 select |
| bus_prdata_uart | 32 | UART→interconnect | UART read data |
| bus_prdata_timer | 32 | Timer→interconnect | Timer read data |
| bus_prdata_gpio | 32 | GPIO→interconnect | GPIO read data |
| bus_prdata_adc | 32 | ADC→interconnect | ADC read data |
| bus_prdata_fir | 32 | FIR→interconnect | FIR read data |
| bus_prdata_qrs | 32 | QRS→interconnect | QRS read data |
| bus_prdata_spi2 | 32 | SPI2→interconnect | SPI2 read data |
| bus_pready_uart | 1 | UART→interconnect | UART access completion |
| bus_pready_timer | 1 | Timer→interconnect | Timer access completion |
| bus_pready_gpio | 1 | GPIO→interconnect | GPIO access completion |
| bus_pready_adc | 1 | ADC→interconnect | ADC access completion |
| bus_pready_fir | 1 | FIR→interconnect | FIR access completion |
| bus_pready_qrs | 1 | QRS→interconnect | QRS access completion |
| bus_pready_spi2 | 1 | SPI2→interconnect | SPI2 access completion |
| bus_prdata | 32 | interconnect→master | Selected read data |
| bus_pready | 1 | interconnect→master | Selected completion |

## 31.3 Canonical Interrupt Signals

| Signal | Source | Meaning |
| --- | --- | --- |
| irq_timer | Timer | Timer STATUS.EXPIRED event enabled by Timer CTRL.IRQ_EN |
| irq_adc_sample | ADC | ADC SAMPLE_RDY event enabled by ADC IRQ_EN[0] |
| irq_adc_overrun | ADC | ADC OVERRUN event enabled by ADC IRQ_EN[1] |
| irq_qrs_beat | QRS | QRS BEAT_DETECTED event enabled by QRS IRQ_EN[0] |
| irq_qrs_arrhythmia | QRS | QRS ARRHYTHMIA_FLAG event enabled by QRS IRQ_EN[1] |
| irq_i | Interrupt OR | irq_timer OR irq_adc_sample OR irq_adc_overrun OR irq_qrs_beat OR irq_qrs_arrhythmia |

This corrects the ambiguity of the earlier generic irq_adc name. Kiro shall use irq_adc_sample and irq_adc_overrun separately, then OR them into irq_i.

## 31.4 Canonical External Interface Signals

| Block | Canonical signals |
| --- | --- |
| ADC SPI | adc_spi_sclk_o, adc_spi_mosi_o, adc_spi_miso_i, adc_spi_cs_n_o |
| UART | uart_tx_o, uart_rx_i |
| GPIO | gpio_led_o, gpio_buzzer_o, gpio_button_i |
| SPI2 | spi2_sclk_o, spi2_mosi_o, spi2_miso_i, spi2_cs_n_o |

## 31.5 Canonical FIR-to-QRS Signals

| Signal | Width | Direction | Meaning |
| --- | --- | --- | --- |
| fir_fifo_wr_valid_o | 1 | FIR→FIR-QRS FIFO | Filtered sample is valid |
| fir_fifo_wr_data_o | 16 | FIR→FIR-QRS FIFO | Signed filtered sample |
| fir_fifo_full_i | 1 | FIR-QRS FIFO→FIR | FIFO cannot accept a write |
| fir_fifo_rd_valid_o | 1 | FIR-QRS FIFO→QRS | FIFO has a valid output sample |
| fir_fifo_rd_data_o | 16 | FIR-QRS FIFO→QRS | Signed filtered sample |
| fir_fifo_rd_ready_i | 1 | QRS→FIR-QRS FIFO | QRS accepts current output sample |

The existing document used fir_fifo_rd_valid_i/fir_fifo_rd_data_i/fir_fifo_rd_ready_o from the QRS point of view. For one universal project-wide convention, the FIFO itself now owns its output names: fir_fifo_rd_valid_o and fir_fifo_rd_data_o; QRS drives fir_fifo_rd_ready_i. Kiro shall use these canonical names in all modules and testbenches.

# 32. Per-Block Control Path and Data Path

For every block, Kiro shall implement two explicitly traceable paths: the CONTROL PATH carries configuration, enables, state transitions, status and interrupts; the DATA PATH carries payload data. A signal shall not silently serve both roles when separate control and data signals can be used.

## 32.1 VeeR EL2 Processor

### Control Path

VeeR EL2
  | bus_paddr, bus_pwrite, bus_pwdata, bus_penable
  v
APB Interconnect
  |
  +--> bus_psel_<block>
  |
  +<-- bus_prdata, bus_pready
  |
  +<-- irq_i
  v
VeeR interrupt/firmware control

The processor initiates all memory-mapped configuration and firmware-mediated sample transfers. It receives irq_i and services the corresponding status registers.

### Data Path

Processor data moves between instruction/data memory and memory-mapped peripherals. ECG sample data moves from ADC FIFO to FIR through firmware bus writes. No CPU transaction is used for the FIR-to-QRS stream.

## 32.2 Instruction Memory

### Control Path

bus_paddr selects an instruction word address. Read enable is implied by an instruction-memory access. No write operation is permitted through the normal firmware path unless the memory wrapper explicitly supports initialization.

### Data Path

Instruction bytes/words are returned from instruction memory to the VeeR instruction interface. The memory does not process ECG data.

VeeR instruction fetch
      |
      +--> bus/instruction address
      |
      v
instruction_memory
      |
      +--> instruction data
      v
VeeR EL2

## 32.3 Data Memory

### Control Path

The processor selects data memory using the 0x0001_0000–0x0001_FFFF address range. Read/write control follows the VeeR memory interface.

### Data Path

Firmware stack, variables, buffers, and software state are stored here. ADC sample values temporarily handled by firmware may pass through CPU registers/memory, but the hardware FIR-to-QRS stream does not.

## 32.4 APB Interconnect

### Control Path

bus_paddr[31:8]
       |
       v
address decoder
       |
       +--> bus_psel_uart
       +--> bus_psel_timer
       +--> bus_psel_gpio
       +--> bus_psel_adc
       +--> bus_psel_fir
       +--> bus_psel_qrs
       +--> bus_psel_spi2
       |
       v
selected bus_pready_<block>
       |
       v
bus_pready

Exactly one peripheral select may be asserted for a valid mapped address. For unmapped accesses all bus_psel_* signals remain low, bus_prdata is zero, and bus_pready is one.

### Data Path

bus_pwdata is forwarded to the selected slave for writes. The selected bus_prdata_<block> is multiplexed onto bus_prdata for reads. No peripheral-to-peripheral data is routed through this interconnect.

## 32.5 Timer

### Control Path

bus_paddr/bus_pwdata
       |
       v
Timer registers
       |
       v
period counter / enable / auto-reload
       |
       +--> timer_expired state
       +--> irq_timer
       +--> 250 Hz acquisition event

The Timer is configured for PERIOD_CYCLES=199999. When enabled, the counter generates one expiry every 200000 clock cycles. With AUTO_RELOAD=1, this creates a 250 Hz event.

### Data Path

The Timer has no ECG payload data path. Its useful output is a control event: the timer expiry starts one ADC acquisition. COUNT is diagnostic state exposed through the bus.

## 32.6 ADC Interface Controller

### Control Path

timer expiry
    |
    v
adc_start_request
    |
    v
ADC SPI FSM
    |
    +--> adc_spi_cs_n_o
    +--> adc_spi_sclk_o
    +--> adc_spi_mosi_o
    |
    +<-- adc_spi_miso_i
    |
    v
sample_complete
    |
    +--> SAMPLE_RDY
    +--> irq_adc_sample
    +--> OVERRUN if FIFO full

The ADC controller must reject a new start while an SPI conversion is active. ADC CTRL.ENABLE gates acquisition. ADC IRQ_EN controls interrupt outputs.

### Data Path

behavioral ADC
      |
adc_spi_miso_i
      |
      v
ADC SPI shift register
      |
12-bit adc_sample_data
      |
      +--> SAMPLE_DATA
      |
      +--> adc_fifo_wr_data
      v
32 x 12 ADC FIFO
      |
      +--> FIFO_DATA
      v
VeeR firmware
      |
FIR SAMPLE_IN write

The ADC controller captures exactly 12 data bits, preserves sample ordering, stores completed samples in the ADC FIFO, and exposes the latest sample through SAMPLE_DATA.

## 32.7 ADC FIFO

### Control Path

Control state consists of adc_fifo_wr_en, adc_fifo_rd_en, read/write pointers, count, full, empty, and overrun detection. These are internal implementation signals; the external architectural status is represented by ADC STATUS.

### Data Path

adc_fifo_wr_data[11:0]
          |
       write pointer
          v
   +----------------+
   | 32 x 12 memory |
   +----------------+
          |
       read pointer
          v
adc_fifo_rd_data[11:0]
          |
       FIFO_DATA

The FIFO is first-in-first-out. A read when empty does not pop or create data. A completed ADC conversion when full is rejected and sets OVERRUN.

## 32.8 FIR Filter

### Control Path

bus_psel_fir + bus_paddr + bus_pwrite + bus_pwdata
                       |
                       v
                  FIR registers
                       |
          +------------+-------------+
          |            |             |
        ENABLE       BYPASS      NOTCH_EN/SEL
          |            |             |
          +------------+-------------+
                       |
                       v
                 FIR control FSM
                       |
                       +--> coefficient selection
                       +--> input acceptance
                       +--> output valid

### Data Path

FIR.SAMPLE_IN[11:0]
        |
        v
center: sample - 2048
        |
        v
16-bit signed input
        |
        +--> sample history x[n]...x[n-30]
        |
        v
31 parallel/serialized MAC operations
        |
        v
40-bit accumulator
        |
        v
arithmetic shift >>> 15
        |
        v
16-bit signed saturation
        |
        v
filtered_sample[15:0]
        |
        +--> SAMPLE_OUT
        |
        +--> fir_fifo_wr_data_o

The FIR data path is completely deterministic. The coefficient address/data registers control the coefficient memory but do not bypass the 31-tap processing structure.

## 32.9 FIR Input FIFO

### Control Path

A bus write to FIR SAMPLE_IN is accepted only if the input FIFO is not full. IN_FIFO_FULL is asserted when count reaches 8. The FIFO preserves firmware write order.

### Data Path

bus_pwdata[11:0]
       |
       v
FIR input FIFO (8 x 16)
       |
       v
centered signed sample
       |
       v
FIR MAC engine

## 32.10 FIR-to-QRS FIFO

### Control Path

fir_fifo_wr_valid_o + !fir_fifo_full_i
                    |
                    v
                 WRITE
                    |
          read/write pointers + count
                    |
fir_fifo_rd_ready_i + fir_fifo_rd_valid_o
                    |
                    v
                  READ

### Data Path

fir_fifo_wr_data_o[15:0]
          |
          v
   32 x 16 FIFO memory
          |
          v
fir_fifo_rd_data_o[15:0]
          |
          v
QRS accelerator

This FIFO is the only continuous hardware connection between FIR and QRS. It is not memory mapped. It shall never be replaced by CPU polling or software copying.

## 32.11 QRS Accelerator

### Control Path

bus_psel_qrs + register writes
             |
             v
       QRS configuration
       ENABLE / RESET_STATS
       THRESHOLD / IRQ_EN
             |
             v
       QRS control FSM
             |
             +--> refractory control
             +--> threshold crossing
             +--> beat event
             +--> RR update
             +--> BPM update
             +--> arrhythmia decision
             |
             +--> irq_qrs_beat
             +--> irq_qrs_arrhythmia

### Data Path

fir_fifo_rd_data_o[15:0]
          |
          v
   derivative engine
          |
          v
      squaring
          |
          v
32-sample moving sum
          |
          v
integrated value
          |
          v
adaptive threshold comparator
          |
          v
     beat event
          |
          +--> RR_INTERVAL
          |
          v
    4-RR averaging
          |
          +--> HEART_RATE
          |
          v
ABS(RR_new - RR_avg)
          |
          v
ARRHYTHMIA_FLAG

QRS asserts fir_fifo_rd_ready_i only when it can accept the current filtered sample. The transfer occurs when fir_fifo_rd_valid_o && fir_fifo_rd_ready_i are both one.

## 32.12 UART

### Control Path

Firmware configures UART CTRL. TX_DATA writes initiate transmission when the transmitter is ready. UART STATUS exposes TX_READY and RX_VALID.

### Data Path

Firmware
  |
bus_pwdata[7:0]
  |
TX_DATA
  |
UART TX shift register
  |
uart_tx_o

Receive data follows the reverse direction through uart_rx_i and RX_DATA. UART is not part of the ECG processing path.

## 32.13 GPIO

### Control Path

Firmware writes DIR and OUT. DIR controls whether a GPIO bit is an output. The application uses OUT[0] for LED and OUT[1] for buzzer.

### Data Path

Firmware --> OUT[1:0] --> gpio_led_o / gpio_buzzer_o
gpio_button_i --> IN[0] --> firmware

## 32.14 SPI2

### Control Path

Firmware
  |
SPI2 CTRL / CS_CTRL / TX_DATA
  |
  v
SPI2 FSM
  |
  +--> spi2_cs_n_o
  +--> spi2_sclk_o
  +--> spi2_mosi_o
  +<-- spi2_miso_i
  |
  +--> TX_BUSY / TX_DONE

### Data Path

TX_DATA[7:0]
      |
      v
SPI2 TX shift register
      |
      v
spi2_mosi_o
      |
external behavioral SPI sink
      |
spi2_miso_i
      |
      v
RX shift register
      |
      v
RX_DATA[7:0]

SPI2 transmits exactly 8 bits per transaction. No OLED or BLE-specific packet interpretation occurs in this block.

## 32.15 Interrupt OR

### Control Path

irq_timer -----------+
irq_adc_sample -------+
irq_adc_overrun ------+
irq_qrs_beat ---------+----> irq_i
irq_qrs_arrhythmia ---+

The interrupt OR is combinational. It contains no state and no bus registers.

### Data Path

There is no payload data path. Interrupt signals are control/status events only.

# 33. Universal End-to-End Signal Trace

Kiro shall be able to trace every ECG sample through the following canonical chain without changing signal names:

clk, rst_n
   |
   +--> timer
   |      |
   |      +--> irq_timer
   |      +--> ADC acquisition start
   |
   +--> adc_interface_controller
   |      |
   |      +--> adc_spi_cs_n_o
   |      +--> adc_spi_sclk_o
   |      +--> adc_spi_mosi_o
   |      +<-- adc_spi_miso_i
   |      |
   |      +--> adc_fifo_wr_data[11:0]
   |      +--> adc_fifo_wr_en
   |
   +--> adc_fifo_32x12
          |
          +--> adc_fifo_rd_data[11:0]
          |
          v
     VeeR firmware
          |
          +--> bus_paddr
          +--> bus_pwrite
          +--> bus_pwdata
          +--> bus_psel_fir
          |
          v
       FIR SAMPLE_IN
          |
          +--> FIR input FIFO
          |
          v
       31-tap FIR
          |
          +--> fir_fifo_wr_valid_o
          +--> fir_fifo_wr_data_o[15:0]
          +<-- fir_fifo_full_i
          |
          v
   fir_qrs_fifo_32x16
          |
          +--> fir_fifo_rd_valid_o
          +--> fir_fifo_rd_data_o[15:0]
          +<-- fir_fifo_rd_ready_i
          |
          v
    QRS accelerator
          |
          +--> HEART_RATE
          +--> RR_INTERVAL
          +--> BEAT_DETECTED
          +--> ARRHYTHMIA_FLAG
          |
          +--> irq_qrs_beat
          +--> irq_qrs_arrhythmia
          |
          v
       irq_i --> VeeR
          |
          +--> UART / GPIO / SPI2

# 34. Universal Naming Rules for Kiro

Use lower_snake_case for all CardioEdge RTL signal names.

Use _i for input ports, _o for output ports, and _io only for true bidirectional signals.

Use clk and rst_n globally; do not create clk_i, clock, reset, rst, reset_n, or alternate names inside CardioEdge.

Use bus_ prefix for all shared memory-mapped bus signals.

Use block-specific prefixes for peripheral signals: adc_, fir_, qrs_, timer_, uart_, gpio_, spi2_.

Use irq_ prefix for interrupt source signals and irq_i for the processor interrupt input.

Use fir_fifo_ prefix for every FIR-to-QRS FIFO signal.

Do not use generic names such as data, valid, ready, enable, status, or irq as top-level interconnect signals when a canonical prefixed name exists.

Internal signals may be named locally, but all module boundary signals shall use the canonical names in this section.

Testbench drivers, monitors, assertions, scoreboards, and waveform documentation shall use the same canonical names.

If a wrapper is necessary to adapt VeeR EL2 native ports to CardioEdge ports, the translation must occur only in the VeeR wrapper; the rest of CardioEdge shall use the canonical names.

Do not duplicate the same physical/logical signal under two different names merely for convenience.

# 35. Block-by-Block Implementation Rule

For every RTL block, Kiro shall create a module-level specification and implementation containing, in this order: (1) purpose, (2) ports using canonical names, (3) control-path description, (4) data-path description, (5) state/FSM behavior, (6) register behavior where applicable, (7) reset behavior, (8) error/recovery behavior, and (9) verification assertions/tests.

This rule applies to the processor wrapper, memories, bus interconnect, UART, Timer, GPIO, ADC controller, ADC FIFO, FIR input FIFO, FIR filter, FIR-QRS FIFO, QRS accelerator, SPI2, and interrupt OR.

The resulting RTL shall therefore be directly traceable from the specification: every control decision has a defined source/destination, and every ECG data word has a defined path from ADC model to QRS output.

# 36. Design Decision Rationale — Why Each Choice Was Made

This section records the engineering reason for every major frozen design decision. The purpose is to prevent Kiro from replacing a deliberate project choice with a different implementation merely because another option is technically possible. The reasons below are design rationale for this RTL-only academic project; they are not claims that the selected value is universally optimal.

Decision hierarchy: the supplied VeeR EL2 reference is authoritative for processor-specific behavior; the CardioEdge specification is authoritative for project architecture; the rationale explains why the CardioEdge values were selected.

## 36.1 Project Scope Decisions

| Decision | Chosen value | Reason |
| --- | --- | --- |
| Implementation scope | RTL + simulation + firmware + verification + synthesis-oriented checks | Keeps the project focused on digital design and allows the entire system to be demonstrated without requiring PCB design, sensor wiring, FPGA bring-up, or physical laboratory equipment. |
| Physical hardware | Not required | The goal is to demonstrate an SoC/RTL architecture. Behavioral models can reproduce the external devices deterministically and make regression repeatable. |
| Technology target | Technology-independent SystemVerilog | Avoids vendor-specific primitives and makes the RTL easier to compile, simulate, review, and potentially synthesize later. |
| Single source of truth | This specification | Prevents architecture drift between Kiro, firmware, RTL, and testbench. |
| No architectural improvisation | MUST/SHALL rules are binding | Prevents an agent from silently adding DMA, changing interfaces, changing algorithms, or selecting incompatible protocols. |

## 36.2 Processor and Bus Decisions

| Decision | Chosen value | Reason |
| --- | --- | --- |
| Processor | Supplied VeeR EL2 | It is the required/reference RISC-V processor for the project; using the supplied core avoids designing a processor and keeps effort on SoC integration. |
| Bus masters | One: VeeR EL2 | A single master eliminates arbitration complexity and makes bus behavior deterministic. It is sufficient because firmware is responsible for ADC-to-FIR transfers. |
| Bus type | APB-style shared memory-mapped bus | The peripherals are low-bandwidth control/status devices. A simple APB-style interface is easier to implement and verify than a high-performance interconnect. |
| No DMA | Explicitly excluded | DMA would add another bus master, descriptors, arbitration, and verification effort. The project already demonstrates hardware acceleration through FIR and QRS, so firmware-mediated ADC transfer is sufficient. |
| Peripheral windows | 256 bytes each | Each peripheral needs only a small register set. Fixed 256-byte windows simplify decoding and make the memory map easy to understand. |
| Address decode | paddr[31:8] for block, paddr[7:2] for register | This matches the 256-byte peripheral windows and 32-bit word alignment while keeping decode logic simple. |

## 36.3 Clock and Reset Decisions

| Decision | Chosen value | Reason |
| --- | --- | --- |
| System clock | 50 MHz | A common, simple digital-system frequency that provides ample timing resolution while making the 250 Hz timer period an exact integer: 50,000,000/250 = 200,000 cycles. |
| ECG sample rate | 250 Hz | High enough for a basic ECG/QRS demonstration and convenient for fixed-point processing. It also makes 4 ms/sample exact and simplifies RR/BPM arithmetic. |
| Reset | Active-low synchronous rst_n | A single synchronous reset avoids asynchronous timing paths and makes simulation/reset behavior deterministic across the RTL hierarchy. |
| Single clock domain | One 50 MHz clk | Avoids CDC logic and associated metastability verification. SPI clocks are generated internally rather than becoming independent RTL clock domains. |

## 36.4 ECG Acquisition Decisions

| Decision | Chosen value | Reason |
| --- | --- | --- |
| ADC model | MCP3201-compatible behavioral model | Provides a concrete, documented SPI ADC behavior while avoiding physical hardware. A real device protocol gives the testbench a meaningful serial transaction to verify. |
| ADC resolution | 12 bits | Matches the selected ADC model and gives sufficient numerical resolution for the educational ECG processing pipeline. |
| ADC channels | 1 | Only one ECG waveform is required; additional channels would add no value to the stated application. |
| ADC SPI mode | Mode 0 | Fixed mode removes unnecessary configurability and matches the selected ADC transaction model. |
| ADC SPI clock | 1 MHz | Well below the selected ADC's capability and simple to generate exactly from 50 MHz using a divide-by-25 half-period. |
| ADC FIFO | 32 × 12 bits | Provides buffering between asynchronous-in-time acquisition events and firmware service without making the FIFO large or difficult to verify. |
| ADC-to-FIR transfer | Firmware-mediated | Keeps the bus/software role visible and avoids DMA. It is also easy to inspect during simulation. |
| No sample overwrite | Mandatory | ECG samples are ordered time-series data. Silent overwrite would make later QRS results non-deterministic and difficult to debug. |

## 36.5 Timer Decisions

| Decision | Chosen value | Reason |
| --- | --- | --- |
| Timer role | 250 Hz acquisition trigger | A dedicated periodic event provides deterministic sample timing and cleanly separates time generation from ADC serial transfer. |
| LOAD | 199,999 | A down-counter starting from 199,999 and including the terminal cycle gives 200,000 clock intervals, exactly 4 ms at 50 MHz. |
| Auto-reload | Supported and used | Continuous ECG acquisition requires repeated triggers; auto-reload avoids firmware having to restart the timer every sample. |
| Timer IRQ | Sticky EXPIRED event | A sticky event prevents firmware from missing a timer expiry during another interrupt service operation. |

## 36.6 FIR Decisions

| Decision | Chosen value | Reason |
| --- | --- | --- |
| Tap count | 31 | 31 taps provide enough impulse-response length for useful ECG band conditioning while remaining small enough for straightforward RTL MAC implementation and verification. |
| Arithmetic | Signed fixed-point | Fixed-point arithmetic is deterministic, synthesizable, resource-predictable, and appropriate for an RTL accelerator. |
| Coefficient format | Q1.15 signed 16-bit | 16-bit coefficients provide practical precision while allowing simple 16×16 multiplication. |
| Accumulator | 40-bit signed | The accumulator must be wider than a single product to prevent overflow while summing 31 signed products. 40 bits gives comfortable headroom for the selected 16×16 operands. |
| Output | 16-bit signed with saturation | A fixed output width simplifies FIFO/register interfaces. Saturation prevents wrap-around distortion when an arithmetic result exceeds the representable output range. |
| Input centering | ADC sample − 2048 | The ADC produces unsigned 0–4095 values. Centering around zero converts the waveform into the signed representation expected by the FIR. |
| Input FIFO | 8 × 16 bits | Provides a small decoupling buffer between firmware writes and FIR processing without creating unnecessary storage. |
| Notch implementation | Coefficient-bank selection | A coefficient-bank approach avoids adding a separate IIR filter, state variables, stability concerns, and extra control logic while still allowing 50/60 Hz rejection behavior. |
| Three banks | BASE, 50 Hz, 60 Hz | The application needs a normal filter plus common power-line rejection choices. Three fixed banks are simple to select and verify. |
| Frozen coefficients | Numerical values fixed in specification | The reference model and RTL must produce identical results. Allowing Kiro to redesign coefficients would make verification results non-reproducible. |

## 36.7 FIR-to-QRS FIFO Decisions

| Decision | Chosen value | Reason |
| --- | --- | --- |
| Depth | 32 | Enough to absorb short bursts of FIR production versus QRS consumption while keeping memory and pointer logic small. |
| Width | 16 bits | Matches the FIR output exactly, avoiding repacking or truncation at the hardware pipeline boundary. |
| Dedicated hardware connection | Yes | QRS is a continuous accelerator and should not depend on CPU polling for every filtered sample. This preserves throughput and demonstrates a meaningful hardware pipeline. |
| Valid/ready interface | Yes | This is a standard simple backpressure mechanism. It allows the FIR to hold data safely when the FIFO is full and allows QRS to consume only when ready. |
| No bus mapping | Yes | The FIFO is an internal streaming resource, not a configuration peripheral. Removing it from the address map simplifies the bus and prevents firmware from interfering with the hardware stream. |
| No overwrite | Yes | Filtered samples must remain chronologically correct for QRS detection. |

## 36.8 QRS Algorithm Decisions

| Decision | Chosen value | Reason |
| --- | --- | --- |
| Derivative | 5-point integer equation | A small fixed stencil captures rapid ECG slope changes while requiring only additions/subtractions and shifts, making it inexpensive in RTL. |
| Derivative scaling | >>> 3 | Division by 8 is implemented as a shift, avoiding a hardware divider. |
| Squaring | Integer square | Makes both positive and negative derivative excursions contribute to the energy measure and is simple to implement. |
| Moving window | 32 samples | At 250 Hz this is 128 ms, a practical short integration window for a basic QRS-energy detector and a power-of-two size that permits division by a simple shift. |
| Integration division | >>> 5 | 32 is a power of two, so averaging needs no divider. |
| Adaptive threshold | running_peak / 4 | A very simple adaptive rule that follows signal magnitude while requiring only a right shift. |
| Minimum threshold | 1 | Prevents zero threshold from causing false threshold crossings during startup. |
| Refractory | 50 samples = 200 ms | Prevents multiple detections from the same cardiac complex and uses an exact sample count that is easy to implement and verify. |
| RR representation | Sample counts | At a fixed 250 Hz sample rate, sample counts directly represent time and avoid a second time-base counter. |
| RR history | 4 intervals | Small enough for simple storage and averaging while providing more stability than relying on only the immediately previous interval. |
| BPM formula | 15000/RR | Because BPM = 60 × Fs / RR and Fs=250 Hz, the constant is exactly 15,000. Integer division avoids floating-point hardware. |
| Arrhythmia rule | ABS(RR_new−RR_avg) > threshold | Simple, deterministic, easy to explain, and sufficient for demonstrating event classification without implementing a clinical arrhythmia classifier. |
| Default RR threshold | 25 samples | At 250 Hz this represents 100 ms. It is a simple conservative demonstration threshold and remains programmable. |
| Startup requirement | 4 valid RR intervals | Prevents arrhythmia comparison against an uninitialized or statistically meaningless history. |
| Peak behavior | No decay algorithm | Avoids another tunable algorithm and keeps the detector deterministic for the academic RTL scope. |

## 36.9 Interrupt Decisions

| Decision | Chosen value | Reason |
| --- | --- | --- |
| Interrupt architecture | Flat OR | Only a few interrupt sources exist. A flat OR is much smaller and easier to verify than a programmable interrupt controller. |
| Interrupt sources | Timer, ADC sample-ready, ADC overrun, QRS beat, QRS arrhythmia | These are the events firmware actually needs to service or report. |
| Sticky event bits | Yes | Firmware may be busy when an event occurs. Sticky status prevents short events from being lost. |
| Clear mechanism | R/W1C | Write-one-to-clear is a simple, common event-register behavior that allows selective clearing without read-modify-write ambiguity. |
| ISR behavior | Inspect all sources | Multiple sources may be asserted simultaneously because irq_i is an OR. The ISR must therefore service all pending causes before returning. |

## 36.10 UART, GPIO and SPI2 Decisions

| Decision | Chosen value | Reason |
| --- | --- | --- |
| UART | Byte-oriented debug/reporting | UART provides a simple way for firmware to expose BPM, RR, status, and diagnostic information in simulation without complicating the ECG datapath. |
| GPIO outputs | LED + buzzer | These give visible/logical representations of normal and alert states without requiring an external protocol. |
| GPIO input | Button | Provides a simple firmware-visible input for demonstrating control/application behavior. |
| SPI2 | Generic byte-oriented SPI master | Demonstrates a second serial interface while avoiding an unnecessary OLED or BLE protocol implementation. |
| SPI2 CPOL/CPHA | Programmable | Unlike the ADC interface, SPI2 has no fixed external device requirement, so limited configurability makes the block more reusable without materially increasing complexity. |
| No OLED protocol | Explicitly excluded | A display protocol would add formatting, command tables, and device-specific behavior unrelated to the RTL architecture being demonstrated. |
| No BLE protocol | Explicitly excluded | BLE requires a complete radio/protocol stack and is far outside the intended RTL-only scope. |

## 36.11 Verification Decisions

| Decision | Chosen value | Reason |
| --- | --- | --- |
| Behavioral external models | ADC + SPI2 sink | Makes external interactions deterministic and allows assertions/scoreboards to verify exact digital behavior. |
| Software reference model | Required | The FIR and QRS algorithms contain fixed-point arithmetic. A bit-accurate reference is the strongest way to detect arithmetic or sequencing errors. |
| Unit testbench per module | Required | Localizes failures and prevents end-to-end debugging from hiding the actual source of an error. |
| End-to-end regression | Required | A design can pass individual block tests while still having incorrect integration, signal naming, register mapping, or data ordering. |
| Waveform checks | Required for serial/control interfaces | SPI, FIFO handshakes, bus phases, and interrupts are timing-sensitive; waveform assertions make those contracts explicit. |
| No physical validation claim | Required | The project is an RTL engineering demonstration and must not be presented as clinically validated ECG hardware. |

# 37. Decision-Change Control

After v1.0 is frozen, Kiro shall not change any decision in Sections 36 or elsewhere in the specification merely to simplify its own implementation. If a genuine contradiction is discovered, Kiro shall stop at the affected block, identify the exact conflicting requirements, and report the conflict. It shall not silently select a replacement value.

If an implementation detail is not specified and does not affect the external contract, Kiro may choose a reasonable RTL structure. Such a choice must not change signal names, register behavior, data widths, timing contracts, FIFO depths, algorithm equations, reset values, or verification expectations.

This rationale section exists specifically so that a reviewer can answer the question 'Why was this value chosen?' without having to infer the reason from the RTL.

# 38. Reviewer Quick-Check: Why This Architecture Is Suitable

The processor handles configuration and supervision rather than every continuous signal-processing operation.

The ADC path is simple and observable: timer → SPI ADC controller → 32-entry FIFO → firmware.

The FIR is a real hardware accelerator with fixed-point arithmetic and coefficient memory.

The FIR-to-QRS FIFO demonstrates hardware streaming and backpressure.

The QRS accelerator demonstrates meaningful digital signal processing without requiring a large or clinically complex algorithm.

The interrupt structure is intentionally simple and fully visible.

Every external device is modeled, so the complete system can be demonstrated in simulation.

All major numerical choices are frozen, making RTL and software reference results reproducible.

The architecture remains small enough to implement and verify as an RTL-only academic project.

# 39. Final Kiro Project Output Contract

The final CardioEdge repository shall be a complete, self-contained RTL project. Kiro shall generate not only RTL but also all documentation, scripts, testbenches, behavioral models, library/package files, register files, firmware, reference models, build infrastructure, run infrastructure, and project status reports. No physical hardware is required.

## 39.1 Required Project Directory

CardioEdge/
├── README.md
├── SPECIFICATION.md
├── LICENSE.txt
├── CHANGELOG.md
├── PROJECT_STATUS.md
├── docs/
├── rtl/
├── registers/
├── firmware/
├── tb/
├── models/
├── reference/
├── scripts/
├── sim/
├── synthesis/
├── data/
└── run

## 39.2 Required File Classes

Documents: README, specification, architecture, block diagrams, control/data paths, signal reference, register reference, firmware guide, verification plan/results, synthesis notes, and design-decision rationale.

Scripts: coefficient generation, register generation, firmware build, RTL compilation, unit regression, end-to-end regression, synthesis check, and cleanup.

RTL: every specified hardware block, common package, top level, and VeeR wrapper.

Testbench: unit testbenches for every stateful/algorithmic block plus complete SoC integration.

Models: MCP3201-compatible ADC model, SPI2 sink model, and ECG stimulus model.

Library/package files: common SystemVerilog package, types/constants, register definitions, verification utilities.

Register files: CSV, JSON, SystemVerilog definitions/package, and human-readable reference.

Run file: executable root `run` that performs the documented default build and regression.

Reference models: frozen coefficients plus bit-accurate FIR/QRS/end-to-end Python models.

## 39.3 Reproducibility

Generated files shall be deterministic.

The same stimulus and source tree shall produce the same numerical results.

Regression shall not silently download unpinned dependencies.

All logs shall be placed under sim/logs.

PROJECT_STATUS.md shall record commands, tool versions where available, tests, pass/fail counts, warnings, and limitations.

# 40. Final Kiro Operating Instructions

Read the complete CardioEdge specification before creating architecture or code.

Locate and inspect the supplied VeeR EL2 Programmer's Reference Manual/reference file. It is authoritative for all VeeR-specific ports, memory interfaces, reset behavior, bus behavior, and processor integration.

Never guess a VeeR port name or interface detail. Adapt the processor only through the wrapper when required.

Use the canonical CardioEdge signal names defined in the specification for every non-VeeR module boundary.

Implement every specified control path, data path, FSM, register, reset behavior, error path, and interrupt behavior.

Use the frozen FIR coefficients and exact integer QRS equations from the specification.

Do not add DMA, another bus master, cache, MMU, another interrupt controller, extra accelerators, OLED/BLE protocols, or physical hardware dependencies.

Do not change sample rate, clock, FIFO sizes, register map, numerical widths, QRS parameters, interrupt semantics, or signal names.

Create the complete directory tree and every required file class.

Build firmware, compile RTL, run unit tests, run integration tests, run end-to-end tests, and run synthesis-oriented checks.

If a requirement is genuinely missing and would affect external behavior, stop and report the exact gap instead of inventing a value.

Internal implementation choices are permitted only when they cannot alter the specified external behavior.

Do not claim completion until the completion checklist passes.

## 40.1 Final Completion Report

Kiro shall produce PROJECT_STATUS.md with the generated directory tree, all RTL modules, testbench/model files, register files, library/package files, firmware, scripts, root run command, tool versions, test results, synthesis results, warnings, known limitations, and an explicit compliance checklist.

# 41. Final Kiro-Ready Implementation Contract

This section is the final implementation-locking layer. It converts the architectural requirements into explicit integration, state-machine, firmware, verification, file-ownership, and acceptance rules. Kiro shall use this section together with the earlier sections; it shall not replace or weaken any earlier normative requirement.

## 41.1 Implementation Order

Kiro shall implement in the following dependency order. A later stage shall not be treated as complete merely because its source files exist.

| Stage | Required work | Exit condition |
| --- | --- | --- |
| 1 | Read supplied VeeR reference and this specification | VeeR integration assumptions documented; no guessed ports |
| 2 | Create cardioedge_pkg and register definitions | All constants/types/register offsets compile |
| 3 | Create memories and APB interconnect | Address decode and bus tests pass |
| 4 | Create Timer/UART/GPIO/SPI2 | Unit tests pass |
| 5 | Create ADC controller + ADC FIFO + MCP3201 model | Serial waveform and FIFO tests pass |
| 6 | Create FIR input FIFO + FIR + frozen coefficients | Bit-accurate FIR tests pass |
| 7 | Create FIR-QRS FIFO | All flow-control tests pass |
| 8 | Create QRS accelerator | Bit-accurate QRS tests pass |
| 9 | Create interrupt OR | All source combinations pass |
| 10 | Create VeeR wrapper/top integration | Processor and peripheral integration compiles |
| 11 | Create firmware/startup/linker/HAL | Firmware builds and initializes system |
| 12 | Create full SoC testbench | Boot/configuration/integration tests pass |
| 13 | Run end-to-end regression | Expected ECG results pass |
| 14 | Run synthesis-oriented checks | Synthesizable RTL hierarchy passes |
| 15 | Generate final documentation/status | All deliverables and checklist complete |

## 41.2 VeeR EL2 Integration Contract

The supplied VeeR EL2 reference is the only authority for the actual processor-side interface. Kiro shall inspect the supplied reference and record the exact VeeR port names, widths, timing, reset expectations, and memory/bus handshake in a document named `docs/veer_integration.md`.

The CardioEdge wrapper shall translate the VeeR interface into the canonical CardioEdge bus names. The wrapper is the only permitted place for naming adaptation.

VeeR EL2 native interface
        |
        v
veer_el2_wrapper.sv
        |
        +--> bus_paddr
        +--> bus_pwrite
        +--> bus_pwdata
        +--> bus_penable
        +<-- bus_prdata
        +<-- bus_pready
        +<-- irq_i
        |
        v
APB interconnect

Kiro shall not alter the VeeR core source unless the supplied reference explicitly requires a project-specific configuration change. Any required VeeR configuration shall be documented.

## 41.3 Top-Level CardioEdge Interface

The top-level module shall expose only the external interfaces required by the project. Exact VeeR-native ports shall remain inside the wrapper as determined from the supplied reference.

| Port | Direction | Width | Purpose |
| --- | --- | --- | --- |
| clk | input | 1 | 50 MHz system clock |
| rst_n | input | 1 | Active-low synchronous reset |
| adc_spi_sclk_o | output | 1 | ADC serial clock |
| adc_spi_mosi_o | output | 1 | ADC serial data output |
| adc_spi_miso_i | input | 1 | ADC serial data input |
| adc_spi_cs_n_o | output | 1 | ADC chip select |
| uart_tx_o | output | 1 | UART transmit |
| uart_rx_i | input | 1 | UART receive |
| gpio_led_o | output | 1 | LED output |
| gpio_buzzer_o | output | 1 | Buzzer output |
| gpio_button_i | input | 1 | Button input |
| spi2_sclk_o | output | 1 | SPI2 serial clock |
| spi2_mosi_o | output | 1 | SPI2 serial output |
| spi2_miso_i | input | 1 | SPI2 serial input |
| spi2_cs_n_o | output | 1 | SPI2 chip select |

No analog input port is required because the ADC is modeled digitally. The ADC model connects to the four ADC SPI signals in the testbench.

## 41.4 FSM Contract

FSM names are normative at the behavioral level. Kiro may choose binary/one-hot encoding, but it shall implement the specified states and transitions.

### ADC SPI FSM

| State | Entry | Actions | Exit |
| --- | --- | --- | --- |
| ADC_IDLE | reset or frame complete | CS high; SCLK idle | start request |
| ADC_START | start accepted | CS low; initialize counters/shift register | start clock |
| ADC_CLOCK_LOW | active transfer | SCLK low; prepare/capture according to Mode 0 timing | half-period elapsed |
| ADC_CLOCK_HIGH | active transfer | SCLK high; capture required data bit | half-period elapsed |
| ADC_COMPLETE | 12 data bits captured | CS high; push sample | one clock |

### FIR FSM

| State | Entry | Actions | Exit |
| --- | --- | --- | --- |
| FIR_IDLE | reset or no input | wait for input FIFO data | data available and enabled |
| FIR_LOAD | input accepted | load/shift sample history; clear accumulator | load complete |
| FIR_MAC | processing | multiply current sample/coefficient and accumulate | tap 30 complete |
| FIR_SCALE | MAC complete | arithmetic shift by 15 | one cycle |
| FIR_SATURATE | scaled result | clamp to -32768..32767 | one cycle |
| FIR_WRITE | output ready | write output FIFO; assert output valid | write accepted |

### QRS FSM

| State | Entry | Actions | Exit |
| --- | --- | --- | --- |
| QRS_IDLE | reset/disabled | hold statistics; no consume | enabled |
| QRS_TRACK | enabled | consume valid filtered samples; derivative/square/integration | threshold crossing or disable |
| QRS_REFRACTORY | beat accepted | count 50 samples; suppress new beats | counter expires |
| QRS_WAIT_BELOW | threshold event | wait for integrated value <= threshold | below threshold |

### SPI2 FSM

| State | Entry | Actions | Exit |
| --- | --- | --- | --- |
| SPI2_IDLE | reset or complete | CS inactive; TX not busy | TX_DATA write while enabled |
| SPI2_ASSERT | start | assert CS; load shift register | start clock |
| SPI2_TRANSFER | active | 8-bit serial transfer using CPOL/CPHA | 8 bits complete |
| SPI2_COMPLETE | 8 bits complete | latch RX_DATA; set TX_DONE; release CS if configured | one cycle |

## 41.5 Firmware Startup Contract

Firmware shall execute this sequence unless the VeeR reference imposes a processor-specific startup requirement that must be adapted without changing the CardioEdge sequence.

Initialize stack/runtime according to the VeeR toolchain/linker requirements.

Disable or mask CardioEdge interrupts during initialization.

Clear GPIO outputs.

Configure Timer LOAD=199999, AUTO_RELOAD=1, IRQ_EN as specified.

Configure ADC for ENABLE=1 and the fixed 1 MHz serial clock.

Drain/clear any stale ADC FIFO state.

Load the exact frozen FIR coefficient table and select BASE.

Configure QRS THRESHOLD=25 and clear QRS statistics.

Enable required QRS interrupts.

Enable required ADC interrupts.

Enable Timer acquisition.

Enable global processor interrupt handling.

Enter the main loop.

On ADC service, drain all available ADC FIFO samples and write them to FIR.SAMPLE_IN until FIR input FIFO is full.

On QRS beat/arrhythmia events, read status, RR_INTERVAL and HEART_RATE as required and clear sticky events with R/W1C.

Use UART/GPIO/SPI2 only for application reporting/control; do not route the continuous FIR-QRS stream through software.

## 41.6 Verification Vectors

| Test ID | Stimulus | Expected result |
| --- | --- | --- |
| V01 | Reset asserted/deasserted | All specified reset values; FIFOs empty; interrupts inactive |
| V02 | Read every mapped register | Correct reset/default value |
| V03 | Write/read writable registers | Value retained subject to defined side effects |
| V04 | Unmapped read/write | prdata=0, pready=1, no side effect |
| V05 | Timer LOAD=199999 at 50 MHz | Expiry every 200000 cycles |
| V06 | ADC injected sample 0xA55 | FIFO_DATA returns 0xA55 |
| V07 | Fill ADC FIFO to 32 | FULL asserted; 33rd completed sample causes OVERRUN/no overwrite |
| V08 | FIR impulse/reference vector | Output matches frozen coefficient reference bit-for-bit |
| V09 | FIR maximum positive/negative stress | Output saturates at 32767/-32768 when required |
| V10 | FIR input FIFO full | Additional write rejected |
| V11 | FIR-QRS FIFO full | FIR output held; no overwrite |
| V12 | FIR-QRS FIFO simultaneous read/write | Count/order remain correct |
| V13 | QRS known synthetic beat spacing | RR_INTERVAL equals exact sample spacing |
| V14 | RR=250 samples | HEART_RATE=60 |
| V15 | RR=125 samples | HEART_RATE=120 |
| V16 | RR=300 samples | HEART_RATE=50 |
| V17 | Four stable RR values | RR average equals integer average |
| V18 | RR deviation <=25 samples | No arrhythmia |
| V19 | RR deviation >25 samples | ARRHYTHMIA_FLAG asserted |
| V20 | Two peaks within <50 samples | Only first beat accepted |
| V21 | Beat and timer simultaneously | Both interrupt causes remain observable |
| V22 | ADC sample-ready and overrun simultaneously | Both causes observable/clearable |
| V23 | SPI2 TX while idle | 8-bit transfer completes |
| V24 | SPI2 TX while busy | Write ignored/no corruption |
| V25 | Processor boot | Firmware reaches configured main loop |
| V26 | End-to-end ECG stimulus | Expected beat/RR/BPM/arrhythmia results |

## 41.7 File-by-File Responsibility

| File/module | Sole responsibility |
| --- | --- |
| cardioedge_top.sv | Top-level integration and external ports |
| veer_el2_wrapper.sv | Only VeeR-to-CardioEdge interface adaptation |
| cardioedge_pkg.sv | Common constants, types, addresses and shared definitions |
| apb_interconnect.sv | Address decode, select generation, read-data mux and completion |
| instruction_memory.sv | Instruction storage/read path |
| data_memory.sv | Data storage/read/write path |
| timer.sv | 250 Hz timing and timer event |
| uart.sv | UART TX/RX |
| gpio.sv | GPIO registers and external GPIOs |
| adc_interface_controller.sv | MCP3201 digital SPI controller |
| adc_fifo.sv | 32×12 ADC buffering |
| fir_input_fifo.sv | 8×16 FIR input buffering |
| fir_filter.sv | 31-tap fixed-point filtering |
| fir_qrs_fifo.sv | 32×16 FIR-to-QRS streaming |
| qrs_accelerator.sv | Derivative, integration, threshold, RR, BPM, arrhythmia |
| spi2_controller.sv | Generic 8-bit SPI2 master |
| interrupt_or.sv | Flat interrupt OR only |
| mcp3201_model.sv | Simulation-only ADC model |
| spi2_sink_model.sv | Simulation-only SPI2 endpoint |
| ecg_stimulus_model.sv | Simulation ECG stimulus |

## 41.8 Final Acceptance Checklist

Every specified block has an RTL source file.

Every required testbench exists and runs.

Every required behavioral model exists.

Every required register file exists and agrees with RTL.

Every required package/library file exists.

Every required firmware file exists and builds.

Every required script exists and is executable where appropriate.

Root `run` performs the complete default regression.

SPECIFICATION.md agrees with the DOCX specification.

Universal signal names are used at every CardioEdge module boundary.

VeeR-specific names are confined to the wrapper/adaptation boundary.

All register offsets, fields, access types, reset values and side effects match.

All control/data paths match the specification.

All FSM behavioral states and transitions match.

All error/recovery paths are implemented.

All interrupts are observable and clearable as specified.

FIR matches the frozen numerical reference.

QRS matches the frozen integer reference.

Firmware sequence matches the specification.

End-to-end ECG regression passes.

Synthesis-oriented checks pass.

No physical hardware is needed.

PROJECT_STATUS.md explicitly lists every test and its result.

## 41.9 Final Rule

Kiro shall treat this document as a deterministic engineering contract, not as a suggestion list. It may optimize internal RTL structure only when the optimization is behaviorally equivalent. Any change that alters a module boundary, canonical signal, register, memory address, timing contract, FIFO behavior, arithmetic result, interrupt behavior, firmware-visible behavior, or required file/deliverable must be rejected unless the specification is explicitly revised.