SCARDIOEDGE SoC

AXI4 SYSTEM SPECIFICATION

v2.0 — IMPLEMENTATION LOCKED

RISC-V Based Real-Time ECG Acquisition, Filtering, Heartbeat Detection and Arrhythmia Alert

AXI4-only memory-mapped system architecture; no APB or AXI4-Lite bridge is used.

| Item | Value |
| --- | --- |
| Source specification | CardioEdge_Specification_FINAL_KIRO_READY.docx |
| Processor reference | RISC-V VeeR EL2 Programmer's Reference Manual, Rev. 1.4 |
| Bus protocol | AMBA AXI4 |
| AXI data width | 64 bits |
| AXI address width | 32 bits |
| System clock | 50 MHz |
| Reset | Active-low synchronous rst_n |
| AXI masters | 2 processor master ports: IFU and LSU |
| AXI slaves | 9: IMEM, DMEM, UART, Timer, GPIO, ADC, FIR, QRS, SPI2 |
| Peripheral register width | 32-bit architectural registers carried over 64-bit AXI4 |
| Version purpose | Complete AXI4 replacement of the bus/interconnect architecture while preserving CardioEdge functional requirements |

# 0. Implementation Rules

This document is the AXI4 revision of the supplied CardioEdge specification. All CardioEdge functional requirements are retained unless an AXI-specific bus rule is explicitly stated here. The memory map, register offsets, FIFO sizes, ECG sample rate, FIR/QRS numerical rules, interrupt sources, external ports, firmware responsibilities, verification intent, and RTL-only scope remain unchanged.

- Use AXI4 throughout the CardioEdge memory-mapped system. No APB interconnect, APB peripheral interface, APB bridge, or AXI4-to-APB bridge is permitted.

- Use the supplied VeeR EL2 Programmer's Reference Manual as the authority for VeeR-native AXI4 ports, widths, reset behavior, clock enables, and handshake semantics.

- The CardioEdge wrapper is the only location where VeeR-specific naming is adapted to the CardioEdge canonical AXI4 interface.

- Use two AXI4 master interfaces from the VeeR core complex: IFU instruction-fetch master and LSU load/store master. This remains one processor/core complex, not two CPUs.

- Use exactly nine AXI4 slave targets in the CardioEdge address map.

- Do not add DMA, a second processor, a programmable interrupt controller, cache, MMU, extra accelerators, OLED/BLE protocols, or other application IP.

- The FIR-to-QRS FIFO remains a dedicated internal valid/ready streaming path and is not memory mapped.

- External ADC and SPI2 endpoint devices remain simulation models; no physical hardware is required.

- All modules reset deterministically with the CardioEdge active-low synchronous reset.

- No requirement shall be silently invented where the source specification is silent. AXI-specific rules introduced by this document are explicitly marked as AXI integration rules.

## 0.1 Source Hierarchy

- For VeeR processor-specific behavior: supplied VeeR EL2 Programmer's Reference Manual.

- For CardioEdge application architecture and numerical behavior: supplied CardioEdge specification.

- For AXI4 system integration details not present in the original APB-oriented specification: this document.

- Where a VeeR-specific signal differs from a CardioEdge signal, the wrapper performs the translation once.

## 0.2 RTL-Only Boundary

The project remains synthesizable SystemVerilog plus firmware, simulation, verification, and synthesis-oriented checks. The ADC is modeled digitally in simulation. UART, GPIO, SPI2, FIR, QRS, memories, and AXI4 interconnect are synthesizable RTL.

# 1. Introduction

CardioEdge is a compact RISC-V-based SoC that acquires digitized ECG samples, performs digital signal conditioning, detects heartbeat/QRS events, calculates heart rate and RR interval, and raises an arrhythmia indication when RR deviation exceeds the configured threshold. The AXI4 revision replaces the original shared APB-style control bus with a complete AXI4 memory-mapped fabric while preserving the control/data separation of the original design.

# 2. Problem Statement

ECG waveforms require signal conditioning before reliable heartbeat detection. Baseline variation and power-line interference can affect peak detection, while processor-only arithmetic creates repeated computation. CardioEdge uses a hardware FIR filter and QRS accelerator, with firmware controlling acquisition, configuration, reporting, and application behavior.

# 3. Objectives

- Integrate the supplied VeeR EL2 RISC-V processor complex.

- Provide 64 KB instruction memory and 64 KB data memory.

- Implement a complete AXI4 memory-mapped system fabric.

- Use VeeR IFU and LSU AXI4 master interfaces.

- Implement AXI4 slave interfaces for UART, Timer, GPIO, ADC Interface Controller, FIR, QRS Accelerator, and SPI2.

- Acquire 12-bit ECG samples at a fixed 250 Hz rate.

- Provide a 32-entry ADC acquisition FIFO.

- Implement a 31-tap signed fixed-point FIR filter.

- Provide optional 50 Hz and 60 Hz notch operation using coefficient-bank selection.

- Provide a 32-entry x 16-bit dedicated FIR-to-QRS FIFO.

- Implement the specified QRS pipeline and deterministic numerical rules.

- Provide BPM, RR interval, beat event, and RR-deviation arrhythmia indication.

- Provide flat interrupt fan-in.

- Implement firmware for initialization, ADC-to-FIR transfer, ISR handling, and reporting.

- Verify the complete SoC using deterministic ECG stimuli and software reference models.

- Keep the design technology-independent and synthesizable.

# 4. System Overview

+----------------------+
                         |      VeeR EL2        |
                         |  IFU AXI4 Master      |
                         |  LSU AXI4 Master      |
                         +----------+-----------+
                                    |
                              AXI4, 64-bit
                                    |
                         +----------v-----------+
                         |   AXI4 INTERCONNECT  |
                         |      2 x 9 fabric    |
                         +--+--+--+--+--+--+---+
                            |  |  |  |  |  |  |
                            v  v  v  v  v  v  v
                           IMEM DMEM UART TMR GPIO ADC FIR QRS SPI2
                                      |
                    ADC -> ADC FIFO -> Firmware -> FIR Input FIFO
                                                   |
                                                   v
                                              31-Tap FIR
                                                   |
                                                   v
                                            FIR-QRS FIFO
                                                   |
                                                   v
                                           QRS Accelerator
                                                   |
                                    BPM / RR / Beat / Arrhythmia

# 5. System Block Definitions

## 5.1 VeeR EL2 RISC-V Core

Processor core complex. The supplied VeeR reference defines separate AXI4 IFU and LSU master system-bus interfaces.

## 5.2 Instruction Memory

64 KB executable memory at 0x0000_0000–0x0000_FFFF. AXI4 slave; normal access is read-only from the processor firmware path.

## 5.3 Data Memory

64 KB runtime memory at 0x0001_0000–0x0001_FFFF. AXI4 slave for firmware stack, variables, and buffers.

## 5.4 AXI4 Interconnect

2-master/9-slave AXI4 fabric. Performs address decode, arbitration, channel routing, response routing, and unmapped-access completion.

## 5.5 UART

Byte-oriented host/debug/reporting peripheral with an AXI4 register interface.

## 5.6 Timer

50 MHz system-clock timer generating the 250 Hz acquisition event.

## 5.7 GPIO

LED, buzzer, and button interface.

## 5.8 ADC Interface Controller

Synthesizable SPI master communicating with the MCP3201-compatible behavioral model and writing 12-bit samples to the ADC FIFO.

## 5.9 FIR Filter

31-tap signed fixed-point filter with 40-bit accumulation, saturation, and selectable coefficient banks.

## 5.10 FIR-to-QRS FIFO

Dedicated 32 x 16-bit valid/ready FIFO outside the AXI4 address map.

## 5.11 QRS Accelerator

Derivative, squaring, 32-sample integration, thresholding, refractory control, RR, BPM, and arrhythmia logic.

## 5.12 Secondary SPI

Generic byte-oriented 8-bit SPI master used for application-result reporting.

## 5.13 External Models

MCP3201-compatible ADC model, SPI2 sink model, and ECG stimulus model are verification-only.

# 6. AXI4 System Architecture

The CardioEdge AXI4 fabric contains two processor-side masters and nine memory-mapped slave targets.

| AXI Master | Source | Allowed accesses | Primary targets |
| --- | --- | --- | --- |
| M0 | VeeR IFU | AXI4 reads only | Instruction Memory |
| M1 | VeeR LSU | AXI4 reads and writes | Data Memory and all peripherals |

| Slave | Block | Base | Limit |
| --- | --- | --- | --- |
| S0 | Instruction Memory | 0x0000_0000 | 0x0000_FFFF |
| S1 | Data Memory | 0x0001_0000 | 0x0001_FFFF |
| S2 | UART | 0x1000_0000 | 0x1000_00FF |
| S3 | Timer | 0x1000_0100 | 0x1000_01FF |
| S4 | GPIO | 0x1000_0200 | 0x1000_02FF |
| S5 | ADC Interface | 0x1000_0300 | 0x1000_03FF |
| S6 | FIR | 0x1000_0400 | 0x1000_04FF |
| S7 | QRS | 0x1000_0500 | 0x1000_05FF |
| S8 | SPI2 | 0x1000_0600 | 0x1000_06FF |

The interconnect shall decode the address using the mapped windows above. The FIR-QRS FIFO has no AXI4 address and is never selected by the interconnect.

# 7. AXI4 Protocol Profile

| Parameter | Requirement |
| --- | --- |
| Protocol | AXI4, not AXI4-Lite |
| Address width | 32 bits |
| Data width | 64 bits |
| Write strobes | 8 bits |
| Read/write IDs | Preserved per VeeR master; interconnect must route IDs with responses |
| Burst support | AXI4 channels shall be structurally supported; VeeR IFU/LSU supplied configuration uses single-beat accesses for the CardioEdge integration unless the VeeR reference specifies otherwise |
| IFU writes | Disabled; VeeR IFU write-valid is hardwired low |
| IFU reads | Single-beat, 64-bit AXI4 read according to VeeR reference |
| LSU reads/writes | 64-bit AXI4 read/write according to VeeR reference |
| Peripheral register accesses | 32-bit architectural register semantics carried on 64-bit AXI4 |
| Reset | No AXI transaction may be falsely completed during reset |

# 8. AXI4 Channel Requirements

Write address: AWADDR, AWVALID, AWREADY, AWID, AWLEN, AWSIZE, AWBURST, AWLOCK,
             AWCACHE, AWPROT, AWQOS, AWREGION
Write data:  WDATA[63:0], WSTRB[7:0], WVALID, WREADY, WLAST
Write resp:  BID, BRESP[1:0], BVALID, BREADY
Read addr:   ARADDR, ARVALID, ARREADY, ARID, ARLEN, ARSIZE, ARBURST, ARLOCK,
             ARCACHE, ARPROT, ARQOS, ARREGION
Read data:   RID, RDATA[63:0], RRESP[1:0], RLAST, RVALID, RREADY

- VALID is held until READY is observed.

- READY may be asserted only when the target can accept the channel transaction.

- Read and write channels are independently handshaked.

- Responses must return the correct transaction ID.

- For single-beat VeeR transactions, RLAST and WLAST are asserted on the accepted data beat.

- The interconnect shall not reorder transactions in a way that violates AXI ordering or VeeR expectations.

- Slave-local register logic may complete an access in one or more cycles; the AXI response is issued only after the corresponding operation has completed.

# 9. AXI4 Interconnect Behavior

- Address decode is performed from the 32-bit byte address.

- Exactly one target is selected for each mapped address.

- M0 IFU accesses to instruction memory route to S0.

- M1 LSU accesses to data memory route to S1.

- M1 LSU accesses in the peripheral range route to S2–S8.

- If M0 and M1 request the same slave concurrently, the interconnect arbitrates while preserving each transaction's ID and response routing.

- Separate read and write channels are handled according to AXI4 handshake rules.

- Unmapped reads complete with RDATA=0 and RRESP=OKAY; unmapped writes complete with BRESP=OKAY and have no side effect. This AXI-specific rule preserves the source requirement that unmapped accesses complete without selecting a slave.

- No target select is generated for reserved addresses.

- No transaction is generated for the internal FIR-QRS FIFO.

# 10. Error and Recovery Policy

| Condition | Required behavior |
| --- | --- |
| Unmapped read | RVALID completes; RDATA=0; RRESP=OKAY; no side effect |
| Unmapped write | BVALID completes; BRESP=OKAY; no side effect |
| ADC FIFO full | Completed conversion is rejected; OVERRUN is set; no overwrite |
| FIR input FIFO full | SAMPLE_IN write has no effect; IN_FIFO_FULL remains asserted |
| FIR-QRS FIFO full | FIR holds output; no overwrite |
| QRS FIFO empty | QRS does not consume a sample |
| SPI2 TX while busy | Write ignored; no queued transfer |
| Sticky event | Event remains set until W1C |

# 11. Interrupt Architecture

irq_timer -----------+
irq_adc_sample -------+
irq_adc_overrun ------+---- OR ----> irq_i --> VeeR EL2 external interrupt input
irq_qrs_beat ---------+
irq_qrs_arrhythmia ---+

- The interrupt OR is combinational and contains no bus registers.

- Timer.STATUS.EXPIRED is enabled by Timer.CTRL.IRQ_EN.

- ADC sample-ready is enabled by ADC.IRQ_EN[0].

- ADC overrun is enabled by ADC.IRQ_EN[1].

- QRS beat is enabled by QRS.IRQ_EN[0].

- QRS arrhythmia is enabled by QRS.IRQ_EN[1].

- The ISR must inspect all relevant status registers because multiple sources may assert simultaneously.

# 12. Memory Map

| Address range | Block | Size | AXI target |
| --- | --- | --- | --- |
| 0x0000_0000–0x0000_FFFF | Instruction Memory | 64 KB | S0 |
| 0x0001_0000–0x0001_FFFF | Data Memory | 64 KB | S1 |
| 0x1000_0000–0x1000_00FF | UART | 256 B | S2 |
| 0x1000_0100–0x1000_01FF | Timer | 256 B | S3 |
| 0x1000_0200–0x1000_02FF | GPIO | 256 B | S4 |
| 0x1000_0300–0x1000_03FF | ADC Interface | 256 B | S5 |
| 0x1000_0400–0x1000_04FF | FIR | 256 B | S6 |
| 0x1000_0500–0x1000_05FF | QRS | 256 B | S7 |
| 0x1000_0600–0x1000_06FF | SPI2 | 256 B | S8 |
| 0x1000_0700–0x1FFF_FFFF | Reserved | — | No target |

# 13. Register Specifications

All architectural registers remain 32 bits. They are transported through the 64-bit AXI4 data bus. For reads, the register value occupies RDATA[31:0] and RDATA[63:32] is zero. For writes, the peripheral consumes WDATA[31:0] with the applicable WSTRB byte lanes. Register accesses are 4-byte aligned.

## 13.1 UART — 0x1000_0000

| Offset | Register | Access | Fields | Reset |
| --- | --- | --- | --- | --- |
| 0x00 | CTRL | R/W | [0] ENABLE; [15:4] BAUD_DIV | 0 |
| 0x04 | STATUS | R | [0] TX_READY; [1] RX_VALID | TX_READY=1 |
| 0x08 | TX_DATA | W | [7:0] TX_BYTE | 0 |
| 0x0C | RX_DATA | R | [7:0] RX_BYTE | 0 |

BAUD_DIV is an integer divider. Exact serial timing is implementation-defined because no physical baud requirement is part of the RTL-only project.

## 13.2 Timer — 0x1000_0100

| Offset | Register | Access | Fields | Reset |
| --- | --- | --- | --- | --- |
| 0x00 | CTRL | R/W | [0] ENABLE; [1] AUTO_RELOAD; [2] IRQ_EN | 0 |
| 0x04 | LOAD | R/W | [31:0] PERIOD_CYCLES | 0 |
| 0x08 | COUNT | R | [31:0] current count | 0 |
| 0x0C | STATUS | R/W1C | [0] EXPIRED | 0 |

LOAD=199999 gives a 200000-cycle period at 50 MHz, equal to 4 ms and 250 Hz. COUNT decrements once per clk while enabled; AUTO_RELOAD reloads LOAD on expiry.

## 13.3 GPIO — 0x1000_0200

| Offset | Register | Access | Fields | Reset |
| --- | --- | --- | --- | --- |
| 0x00 | DIR | R/W | [2:0] LED/Buzzer/Button direction; 1=output | 0 |
| 0x04 | OUT | R/W | [1:0] LED/Buzzer outputs | 0 |
| 0x08 | IN | R | [0] Button input | 0 |

## 13.4 ADC — 0x1000_0300

| Offset | Register | Access | Fields | Reset |
| --- | --- | --- | --- | --- |
| 0x00 | CTRL | R/W | [0] ENABLE; [7:4] CLK_DIV | 0 |
| 0x04 | STATUS | R/W1C | [0] SAMPLE_RDY; [1] FIFO_FULL; [2] FIFO_EMPTY; [3] OVERRUN | FIFO_EMPTY=1 |
| 0x08 | SAMPLE_DATA | R | [11:0] latest sample | 0 |
| 0x0C | FIFO_DATA | R | [11:0] next sample; read pops | 0 |
| 0x10 | IRQ_EN | R/W | [0] SAMPLE_RDY_IE; [1] OVERRUN_IE | 0 |

## 13.5 FIR — 0x1000_0400

| Offset | Register | Access | Fields | Reset |
| --- | --- | --- | --- | --- |
| 0x00 | CTRL | R/W | [0] ENABLE; [1] BYPASS; [2] NOTCH_EN; [3] NOTCH_SEL | 0 |
| 0x04 | STATUS | R/W1C | [0] OUT_VALID; [1] IN_FIFO_FULL | 0 |
| 0x08 | COEFF_ADDR | R/W | [4:0] coefficient index 0–30 | 0 |
| 0x0C | COEFF_DATA | R/W | [15:0] signed Q1.15 coefficient | 0 |
| 0x10 | SAMPLE_IN | W | [11:0] ADC sample | 0 |
| 0x14 | SAMPLE_OUT | R | [15:0] filtered sample | 0 |

## 13.6 QRS — 0x1000_0500

| Offset | Register | Access | Fields | Reset |
| --- | --- | --- | --- | --- |
| 0x00 | CTRL | R/W | [0] ENABLE; [1] RESET_STATS | 0 |
| 0x04 | STATUS | R/W1C | [0] BEAT_DETECTED; [1] ARRHYTHMIA_FLAG | 0 |
| 0x08 | HEART_RATE | R | [15:0] BPM | 0 |
| 0x0C | RR_INTERVAL | R | [31:0] RR samples | 0 |
| 0x10 | THRESHOLD | R/W | [31:0] RR deviation threshold in samples | 0 |
| 0x14 | IRQ_EN | R/W | [0] BEAT_IE; [1] ARRHYTHMIA_IE | 0 |

## 13.7 SPI2 — 0x1000_0600

| Offset | Register | Access | Fields | Reset |
| --- | --- | --- | --- | --- |
| 0x00 | CTRL | R/W | [0] ENABLE; [3:1] CLK_DIV; [4] CPOL; [5] CPHA | 0 |
| 0x04 | STATUS | R | [0] TX_BUSY; [1] TX_DONE | 0 |
| 0x08 | TX_DATA | W | [7:0] byte; starts transfer | 0 |
| 0x0C | RX_DATA | R | [7:0] received byte | 0 |
| 0x10 | CS_CTRL | R/W | [0] CS_ASSERT | 0 |

# 14. ADC Interface Controller

| Parameter | Value |
| --- | --- |
| Resolution | 12 bits |
| Channels | 1 |
| SPI mode | Mode 0 |
| System clock | 50 MHz |
| ECG sample rate | 250 Hz |
| Acquisition FIFO | 32 x 12 bits |
| ADC serial clock | 1 MHz fixed target; CLK_DIV implements the selected division |

- Timer expiry initiates one ADC conversion request.

- A new conversion is not started while a previous SPI conversion is active.

- The controller asserts CS low, generates SPI clocks, captures serial conversion data on the required edge, assembles 12 bits, deasserts CS, updates SAMPLE_DATA, pushes the sample to the FIFO, and sets SAMPLE_RDY.

- If conversion completes while the FIFO is full, the sample is rejected and OVERRUN is set.

- No sample overwrite is permitted.

# 15. FIR Filter IP

- 31 taps, signed Q1.15 coefficients.

- Input is centered: input_s = sign_extend(ADC_sample) - 2048.

- Each product is 16 x 16 -> 32-bit signed.

- Accumulator is 40-bit signed.

- scaled = arithmetic_shift(accumulator, 15).

- Output is saturated to signed 16-bit range -32768 through +32767.

- NOTCH_EN=0 selects BASE. NOTCH_EN=1 selects 50 Hz or 60 Hz coefficient bank using NOTCH_SEL.

- FIR input FIFO is 8 x 16 bits.

- A SAMPLE_IN write is accepted only when the input FIFO has space.

- SAMPLE_OUT holds the latest completed filtered sample and OUT_VALID is sticky until W1C.

# 16. FIR-to-QRS Dedicated FIFO

| Parameter | Value |
| --- | --- |
| Depth | 32 entries |
| Width | 16 bits |
| Storage | 512 bits |
| Interface | valid/ready |
| Bus mapped | No |
| Overwrite | Forbidden |

- Write occurs when fir_fifo_wr_valid_o is asserted and fir_fifo_full_i is low.

- Read occurs when fir_fifo_rd_valid_o and fir_fifo_rd_ready_i are both high.

- Simultaneous read/write is supported and leaves occupancy unchanged.

- When full, FIR holds its output until accepted.

- When empty, valid is low and QRS must not consume.

- Reset clears pointers and count.

# 17. QRS Accelerator

- Consumes filtered 16-bit samples only through the dedicated FIR-QRS FIFO.

- Uses the specified 5-point integer derivative with >>>3 scaling.

- Squares the derivative as an integer.

- Uses a 32-sample moving-window integration.

- Integration average uses >>>5.

- Adaptive threshold is running_peak / 4 with minimum threshold 1.

- Uses a 50-sample refractory period.

- RR intervals are represented as sample counts.

- Maintains four valid RR intervals for the running average.

- BPM = 15000 / RR_INTERVAL using integer division.

- Arrhythmia is asserted when ABS(RR_new - RR_AVG) > THRESHOLD after four valid RR intervals exist.

- Default firmware threshold is 25 samples.

- HEART_RATE remains zero until the first valid RR interval.

- RESET_STATS clears beat history, previous-sample history, RR history, running peak, integration state, BPM, and event state while leaving ENABLE unchanged.

# 18. Secondary SPI Controller

- Generic byte-oriented 8-bit SPI master.

- Programmable CPOL and CPHA.

- TX_DATA write while idle starts one exactly 8-bit transaction.

- TX_BUSY prevents a second transfer; writes while busy are ignored.

- Received byte is captured in RX_DATA and TX_DONE is set after completion.

- No timeout or retry engine.

- No OLED or BLE packet interpretation.

# 19. Firmware Requirements

- Initialize stack/runtime according to the VeeR toolchain/linker requirements.

- Disable or mask CardioEdge interrupts during initialization.

- Clear GPIO outputs.

- Configure Timer LOAD=199999, AUTO_RELOAD=1, IRQ_EN as required.

- Configure ADC ENABLE=1 and the fixed 1 MHz serial clock.

- Drain/clear stale ADC FIFO state.

- Load the exact frozen FIR coefficient table and select BASE.

- Configure QRS THRESHOLD=25 and clear QRS statistics.

- Enable QRS and ADC interrupts as required.

- Enable Timer acquisition.

- Enable global processor interrupt handling.

- On ADC service, drain ADC FIFO_DATA while non-empty and write samples to FIR.SAMPLE_IN until the FIR input FIFO is full.

- On QRS beat/arrhythmia, read STATUS, RR_INTERVAL, and HEART_RATE as required and clear sticky events using W1C.

- Use UART/GPIO/SPI2 for reporting/control only; never route the continuous FIR-QRS stream through software.

# 20. External Output and Alert Behavior

- GPIO provides immediate simulation-visible LED and buzzer indication.

- UART and SPI2 provide serial observation of application state.

- On ARRHYTHMIA_FLAG, firmware sets LED and buzzer outputs and transmits an application status message/byte through UART/SPI2.

- The exact human-readable UART string is not hardware normative.

# 21. Verification Plan

tb_cardioedge
├── clock/reset
├── VeeR firmware/program image
├── AXI4 master/slave monitors
├── AXI4 protocol assertions
├── MCP3201 behavioral ADC model
├── ECG stimulus generator/file reader
├── SPI2 sink model
├── UART monitor
├── GPIO monitor
├── scoreboards
├── reference-model checker
└── end-to-end checks

- AXI4 AW/W/B and AR/R handshake and response-ID checks.

- AXI4 address decode and unmapped accesses.

- Register reset/read/write/side effects.

- Timer countdown, reload, expiry, W1C.

- GPIO direction and output/input.

- ADC SPI transaction, sample assembly, FIFO, overrun.

- FIR coefficient loading, input FIFO, arithmetic, saturation, output.

- FIR-QRS FIFO empty/full/simultaneous read-write/backpressure.

- QRS derivative, square, integration, threshold, refractory, RR, BPM, arrhythmia.

- SPI2 busy/done/receive/chip-select.

- Interrupt OR and sticky-event clearing.

- CPU can access every mapped peripheral.

- Only one AXI target is selected for each mapped address.

- Unmapped AXI access completes without side effect.

- Timer produces 250 Hz acquisition events.

- ADC samples reach firmware in order.

- Firmware transfers samples to FIR in order.

- FIR output reaches QRS through the dedicated FIFO without AXI mediation.

- FIFO backpressure never overwrites samples.

- QRS events reach irq_i.

- Firmware reads BPM/RR and controls GPIO/UART/SPI2.

- Reset returns the complete system to a known state.

# 22. Hardware / Synthesis Requirements

- Use synthesizable SystemVerilog.

- Use explicit signed/unsigned declarations and casts.

- Use explicit arithmetic widths.

- No floating point in synthesizable signal-processing logic.

- No # delays in synthesizable RTL.

- No testbench-only file I/O in synthesizable modules.

- Use synchronous active-low reset.

- Do not infer memories/registers with ambiguous read/write semantics.

- Avoid vendor-specific primitives unless a separate target wrapper is added outside CardioEdge.

- Nominal design clock is 50 MHz.

- No clock-domain crossing is required in v1.0; SPI clocks are generated internally.

- The complete AXI4 hierarchy must compile and pass synthesis-oriented checks.

# 23. Deliverables

- Complete CardioEdge SystemVerilog RTL.

- AXI4 top-level integration.

- VeeR EL2 AXI4 integration wrapper.

- Instruction and data memories.

- 2-master/9-slave AXI4 interconnect.

- AXI4 UART, Timer, GPIO, ADC, FIR, QRS, and SPI2 interfaces.

- ADC controller and 32 x 12 ADC FIFO.

- 31-tap FIR and coefficient storage.

- 32 x 16 dedicated FIR-QRS FIFO.

- QRS/Heartbeat Detection Accelerator.

- Interrupt OR logic.

- Firmware, linker/startup support, and build configuration.

- MCP3201-compatible ADC behavioral model.

- SPI2 sink/testbench model.

- ECG stimulus/reference model.

- Unit, subsystem, and full-SoC verification.

- AXI4 assertions, scoreboards, and monitors.

- Synthesis-oriented scripts/reports.

- Final documentation and demonstration logs/waveforms.

# 24. Frozen Design Decisions

| Decision | Value |
| --- | --- |
| Bus protocol | AXI4 only |
| AXI data width | 64 bits |
| AXI address width | 32 bits |
| AXI masters | VeeR IFU and VeeR LSU |
| AXI slaves | 9 mapped targets |
| Register width | 32-bit semantics on 64-bit AXI |
| ECG sample rate | 250 Hz |
| Timer period | 200,000 system-clock cycles |
| Timer LOAD | 199,999 |
| System clock | 50 MHz |
| ADC resolution | 12 bits |
| ADC FIFO | 32 x 12 |
| FIR input FIFO | 8 x 16 |
| FIR-QRS FIFO | 32 x 16 |
| FIR taps | 31 |
| FIR coefficients | Signed Q1.15, frozen tables |
| FIR accumulator | 40-bit signed |
| QRS derivative | 5-point integer, >>>3 |
| QRS integration | 32 samples, >>>5 |
| QRS refractory | 50 samples |
| RR history | 4 intervals |
| BPM | 15000/RR, integer division |
| Default RR threshold | 25 samples |
| QRS startup | 4 valid RR intervals |
| Interrupts | Flat OR |
| Sticky events | R/W1C |
| Physical hardware | Not required |

# 25. Implementation Order

- Inspect the supplied VeeR EL2 reference and record exact AXI4 port names and widths.

- Create cardioedge_pkg and AXI4/register constants.

- Create instruction/data memories and AXI4 interconnect.

- Create AXI4 Timer/UART/GPIO/SPI2 slaves.

- Create AXI4 ADC controller, ADC FIFO, and MCP3201 model.

- Create FIR input FIFO, FIR, and frozen coefficient banks.

- Create dedicated FIR-QRS FIFO.

- Create QRS accelerator.

- Create interrupt OR.

- Create VeeR AXI4 wrapper and top integration.

- Create firmware/startup/linker/HAL.

- Create full SoC AXI4 testbench.

- Run end-to-end ECG regression.

- Run synthesis-oriented checks.

- Generate final documentation/status.

# 26. Final Implementation Contract

A design that compiles but changes the architecture, register map, data path, numerical rules, interrupt semantics, FIFO behavior, or firmware-visible behavior is not compliant. The AXI4 fabric is the only system memory-mapped bus. No APB-compatible interface shall remain at a CardioEdge module boundary.

# 27. Frozen Numerical Tables and Exact Interface Rules

## 27.1 FIR Coefficient Quantization

The supplied specification freezes three 31-tap coefficient banks. Quantization is signed Q1.15 using integer = round(real_coefficient × 32768), followed by signed 16-bit clipping. The BASE table below is reproduced from the source specification.

| Tap | Integer | Hex | Approx. real |
| --- | --- | --- | --- |
| 0 | 22 | 0x0016 | 0.000671387 |
| 1 | 55 | 0x0037 | 0.001678467 |
| 2 | 31 | 0x001F | 0.000946045 |
| 3 | -92 | 0xFFA4 | -0.002807617 |
| 4 | -250 | 0xFF06 | -0.007629395 |
| 5 | -230 | 0xFF1A | -0.007019043 |
| 6 | 117 | 0x0075 | 0.003570557 |
| 7 | 565 | 0x0235 | 0.017242432 |
| 8 | 522 | 0x020A | 0.015930176 |
| 9 | -383 | 0xFE81 | -0.011688232 |
| 10 | -1625 | 0xF9A7 | -0.049591064 |
| 11 | -1811 | 0xF8ED | -0.055267334 |
| 12 | 277 | 0x0115 | 0.008453369 |
| 13 | 4397 | 0x112D | 0.134185791 |
| 14 | 8571 | 0x217B | 0.261566162 |
| 15 | 10334 | 0x285E | 0.315368652 |
| 16 | 8571 | 0x217B | 0.261566162 |
| 17 | 4397 | 0x112D | 0.134185791 |
| 18 | 277 | 0x0115 | 0.008453369 |
| 19 | -1811 | 0xF8ED | -0.055267334 |
| 20 | -1625 | 0xF9A7 | -0.049591064 |
| 21 | -383 | 0xFE81 | -0.011688232 |
| 22 | 522 | 0x020A | 0.015930176 |
| 23 | 565 | 0x0235 | 0.017242432 |
| 24 | 117 | 0x0075 | 0.003570557 |
| 25 | -230 | 0xFF1A | -0.007019043 |
| 26 | -250 | 0xFF06 | -0.007629395 |
| 27 | -92 | 0xFFA4 | -0.002807617 |
| 28 | 31 | 0x001F | 0.000946045 |
| 29 | 55 | 0x0037 | 0.001678467 |
| 30 | 22 | 0x0016 | 0.000671387 |

## 27.2 Coefficient Bank Selection

- COEFF_ADDR[4:0] selects tap 0 through tap 30.

- COEFF_DATA[15:0] stores signed two's-complement Q1.15 coefficient data.

- BASE is the active bank after reset.

- NOTCH_EN=0 selects BASE.

- NOTCH_EN=1 and NOTCH_SEL=0 selects NOTCH_50HZ.

- NOTCH_EN=1 and NOTCH_SEL=1 selects NOTCH_60HZ.

- The complete locked NOTCH_50HZ and NOTCH_60HZ values remain source-controlled exactly as supplied in the original specification; no coefficient regeneration is permitted.

NOTE: The searchable source extract available while generating this document exposed the complete BASE table and portions of the notch tables, but not a reliable complete machine-readable copy of both notch tables. Therefore this document intentionally does not invent missing notch values. The implementation must copy the exact original source-controlled notch tables into the coefficient package/reference files before regression.

# 28. AXI4 Requirement Traceability Matrix

| ID | Requirement | Implementation | Verification |
| --- | --- | --- | --- |
| AXI-01 | VeeR IFU AXI4 master | veer_el2_axi_wrapper | AXI IFU read test |
| AXI-02 | VeeR LSU AXI4 master | veer_el2_axi_wrapper | AXI LSU read/write test |
| AXI-03 | 2 x 9 AXI4 fabric | axi4_interconnect | Routing/arbitration tests |
| AXI-04 | 64-bit AXI data path | AXI4 interconnect/slaves | Width/byte-strobe tests |
| AXI-05 | Unmapped access completes without side effect | AXI4 interconnect | Unmapped read/write test |
| ARCH-01 | 64 KB IMEM/DMEM | Memory slaves | Address/firmware tests |
| ARCH-02 | 250 Hz acquisition | Timer/ADC | Timer/ADC test |
| ADC-01 | MCP3201-compatible ADC | ADC RTL + model | SPI waveform checker |
| ADC-02 | 32 x 12 ADC FIFO | ADC FIFO | FIFO tests |
| FIR-01 | 31 taps | FIR RTL | Reference test |
| FIR-02 | Frozen Q1.15 coefficients | Coefficient banks | Readback/reference |
| FIR-03 | 40-bit MAC | FIR RTL | Bit-accurate test |
| FIFO-01 | 32 x 16 FIR-QRS FIFO | Dedicated FIFO | Flow-control tests |
| QRS-01 | 5-point derivative | QRS RTL | Reference comparison |
| QRS-02 | 32-sample integration | QRS RTL | Reference comparison |
| QRS-03 | 50-sample refractory | QRS RTL | Double-peak test |
| QRS-04 | Four-RR average | QRS RTL | RR sequence test |
| QRS-05 | BPM=15000/RR | QRS RTL | Known-RR test |
| QRS-06 | RR deviation arrhythmia | QRS RTL | Abnormal-RR test |
| IRQ-01 | Flat interrupt OR | interrupt_or | Multiple-source test |
| REG-01 | Exact register map | AXI4 peripherals | Register regression |
| FW-01 | ADC FIFO to FIR | Firmware | End-to-end |
| FW-02 | ISR source handling | Firmware | Interrupt regression |
| RTL-01 | Synthesizable generic RTL | All RTL | Compile/synthesis |
| SIM-01 | No physical hardware dependency | Testbench models | Full regression |

# 29. Completion Criteria

- All specified RTL modules and top-level ports exist.

- All AXI4 master/slave interfaces compile with no unresolved ports.

- All register offsets, access types, reset values, and side effects match.

- All AXI4 handshake, response, ID, and unmapped-access tests pass.

- All FIFO full/empty/wrap/simultaneous read-write cases pass.

- ADC controller and behavioral model agree on the 12-bit sample sequence.

- FIR output matches the frozen coefficient reference.

- QRS outputs match the integer reference model.

- Interrupt events are latched, routed, serviced, and cleared correctly.

- Firmware boots and initializes the SoC.

- Firmware moves ADC samples to FIR and handles QRS events.

- End-to-end ECG stimulus produces deterministic expected beat/RR/BPM/arrhythmia results.

- No unexplained X/Z values remain on architecturally relevant signals after reset.

- Synthesis/compile checks pass for the complete synthesizable hierarchy.

- No physical hardware is required for regression.

# 30. External Device / Simulation Boundary

- The MCP3201-compatible ADC is a simulation model connected to the four ADC SPI signals.

- SPI2 external destination is a simulation sink model.

- ECG stimulus may be generated from a deterministic file/model.

- No physical ADC, ECG electrodes, FPGA pins, OLED, BLE module, PCB, or electrical-level design is required.

# 31. Universal AXI4 Signal Naming

The original universal CardioEdge naming convention is retained for non-AXI signals. Shared memory-mapped bus signals are now AXI4-prefixed and channel-specific.

| Signal group | Width | Direction | Meaning |
| --- | --- | --- | --- |
| clk | 1 | input | 50 MHz system clock |
| rst_n | 1 | input | Active-low synchronous reset |
| bus_aw* | AXI-dependent | master/interconnect/slave | AXI4 write-address channel |
| bus_w* | AXI-dependent | master/interconnect/slave | AXI4 write-data channel |
| bus_b* | AXI-dependent | interconnect/master | AXI4 write response |
| bus_ar* | AXI-dependent | master/interconnect/slave | AXI4 read-address channel |
| bus_r* | AXI-dependent | interconnect/master | AXI4 read data/response |
| irq_i | 1 | input to VeeR | Combined CardioEdge interrupt |

Canonical block-specific external signals remain: adc_spi_*, uart_*, gpio_*, spi2_*, fir_fifo_*, irq_*.

# 32. Per-Block AXI4 Control/Data Paths

VeeR IFU AXI4
   |
   +----> axi4_interconnect ----> instruction_memory_axi

VeeR LSU AXI4
   |
   +----> axi4_interconnect ----> data_memory_axi
                              +--> uart_axi
                              +--> timer_axi
                              +--> gpio_axi
                              +--> adc_interface_axi
                              +--> fir_axi
                              +--> qrs_axi
                              +--> spi2_axi

ADC --> adc_fifo_32x12 --> VeeR firmware
                         |
                         +--> AXI4 write --> FIR.SAMPLE_IN
FIR --> fir_qrs_fifo_32x16 --> QRS
QRS --> irq_qrs_* --> interrupt_or --> irq_i

# 33. Universal End-to-End Signal Trace

clk, rst_n
 |
 +--> timer
 |     +--> irq_timer
 |     +--> ADC acquisition start
 |
 +--> adc_interface_controller
 |     +--> adc_spi_cs_n_o
 |     +--> adc_spi_sclk_o
 |     +--> adc_spi_mosi_o
 |     +<-- adc_spi_miso_i
 |     +--> adc_fifo_wr_data[11:0]
 |     +--> adc_fifo_wr_en
 |
 +--> adc_fifo_32x12
 |
 +--> VeeR LSU AXI4
 |     +--> AXI4 read of ADC.FIFO_DATA
 |     +--> AXI4 write of FIR.SAMPLE_IN
 |
 +--> FIR input FIFO
 |
 +--> 31-tap FIR
 |     +--> fir_fifo_wr_valid_o
 |     +--> fir_fifo_wr_data_o[15:0]
 |     +<-- fir_fifo_full_i
 |
 +--> fir_qrs_fifo_32x16
 |     +--> fir_fifo_rd_valid_o
 |     +--> fir_fifo_rd_data_o[15:0]
 |     +<-- fir_fifo_rd_ready_i
 |
 +--> QRS accelerator
       +--> HEART_RATE
       +--> RR_INTERVAL
       +--> BEAT_DETECTED
       +--> ARRHYTHMIA_FLAG
       +--> irq_qrs_beat
       +--> irq_qrs_arrhythmia
 |
 +--> interrupt_or --> irq_i --> VeeR

# 34. Universal Naming Rules

- Use lower_snake_case for CardioEdge RTL names.

- Use _i for input ports, _o for output ports, and _io only for true bidirectional signals.

- Use clk and rst_n globally.

- Use bus_ prefix for CardioEdge AXI4 channel groups.

- Use block-specific prefixes for adc_, fir_, qrs_, timer_, uart_, gpio_, spi2_.

- Use irq_ prefix for interrupt source signals and irq_i for the processor interrupt input.

- Use fir_fifo_ prefix for every FIR-to-QRS FIFO signal.

- Do not create APB signal aliases such as bus_paddr, bus_pwrite, bus_penable, bus_psel_* or bus_pready_*.

- VeeR-specific names are confined to the wrapper/adaptation boundary.

- Testbench drivers, monitors, assertions, scoreboards, and waveform documentation use the same canonical names.

# 35. Block-by-Block Implementation Rule

For every RTL block, the module-level specification shall define: (1) purpose, (2) ports, (3) control path, (4) data path, (5) state/FSM behavior, (6) AXI4 register behavior where applicable, (7) reset behavior, (8) error/recovery behavior, and (9) verification assertions/tests.

# 36. AXI4 Design Decision Rationale

| Decision | Chosen value | Reason |
| --- | --- | --- |
| System bus | AXI4 | Matches the supplied VeeR EL2 core-complex system-bus capability and supports the complete memory/peripheral fabric. |
| Data width | 64-bit | VeeR EL2 AXI4 system-bus interfaces use 64-bit data channels. |
| Master structure | IFU + LSU | VeeR exposes separate instruction-fetch and load/store AXI4 masters; both belong to the same core complex. |
| Interconnect | 2 x 9 | Provides explicit routing from both VeeR AXI4 masters to the two memories and seven memory-mapped peripherals. |
| Peripheral interface | AXI4 | Avoids a protocol bridge and makes the CardioEdge memory-mapped fabric AXI4-only. |
| Register semantics | 32-bit on 64-bit AXI | Preserves the original firmware-visible register map while matching the processor bus width. |
| Unmapped response | Complete without side effect | Preserves the source specification's unmapped-access behavior. |
| FIR-QRS bus mapping | No AXI mapping | Continuous signal-processing stream must remain hardware-mediated and independent of processor traffic. |
| Flat interrupt OR | Retained | Only a small number of event sources exist; simple and directly observable. |

## 36.1 Clock and Reset

One 50 MHz synchronous clock is used. rst_n is active-low synchronous reset. The VeeR reference requires the core clock to be stable before core reset deassertion; CardioEdge keeps all system bus logic in the same clock domain.

## 36.2 AXI4 Width and Register Adaptation

The processor-facing bus is 64 bits, while CardioEdge registers remain 32 bits. Peripheral slaves decode the low address bits and expose the register value in RDATA[31:0]. Writes use WDATA[31:0] and the relevant lower WSTRB lanes. The upper data half is ignored for register semantics.

## 36.3 AXI4 Arbitration

The interconnect must arbitrate concurrent requests from IFU and LSU when they target the same slave. The arbitration policy may be fixed-priority or round-robin internally provided it does not starve either master and preserves AXI ordering/response correctness. No additional external master is created.

# 37. Decision-Change Control

After this AXI4 v2.0 specification is frozen, implementation shall not change memory addresses, register offsets, FIFO depths, numerical rules, interrupt semantics, external ports, AXI data width, master/slave topology, or firmware-visible behavior without an explicit specification revision.

# 38. Reviewer Quick-Check

- VeeR EL2 is the only processor complex.

- VeeR IFU and LSU are the only CardioEdge AXI4 master ports.

- The complete memory-mapped fabric is AXI4.

- There are exactly nine AXI4 slave targets.

- There is no APB or AXI4-Lite bridge.

- The address map is unchanged.

- ADC acquisition remains timer-driven at 250 Hz.

- ADC-to-FIR remains firmware-mediated.

- FIR-to-QRS remains hardware-streamed through a 32 x 16 FIFO.

- QRS produces beat, RR, BPM, and arrhythmia information.

- Interrupts remain a flat OR.

- All external devices remain simulation models.

# 39. Final Project Output Contract

CardioEdge/
├── README.md
├── SPECIFICATION.md
├── LICENSE.txt
├── CHANGELOG.md
├── PROJECT_STATUS.md
├── docs/
│   ├── architecture.md
│   ├── axi4_interconnect.md
│   ├── veer_integration.md
│   ├── register_reference.md
│   ├── firmware_guide.md
│   ├── verification_plan.md
│   └── design_decisions.md
├── rtl/
│   ├── cardioedge_top.sv
│   ├── veer_el2_axi_wrapper.sv
│   ├── cardioedge_pkg.sv
│   ├── axi4_interconnect.sv
│   ├── instruction_memory_axi.sv
│   ├── data_memory_axi.sv
│   ├── uart_axi.sv
│   ├── timer_axi.sv
│   ├── gpio_axi.sv
│   ├── adc_interface_controller_axi.sv
│   ├── adc_fifo_32x12.sv
│   ├── fir_input_fifo.sv
│   ├── fir_filter_31tap.sv
│   ├── fir_qrs_fifo_32x16.sv
│   ├── qrs_heartbeat_accelerator.sv
│   ├── spi2_controller_axi.sv
│   └── interrupt_or.sv
├── registers/
├── firmware/
├── tb/
├── models/
├── reference/
├── scripts/
├── sim/
├── synthesis/
├── data/
└── run

# 40. Final AXI4 Operating Instructions

- Read this specification and the supplied VeeR EL2 reference before coding.

- Never guess VeeR native port names; use the exact names and widths from the reference.

- Do not retain APB signal names or APB protocol logic in the CardioEdge module hierarchy.

- Implement the AXI4 interconnect before integrating the peripherals.

- Verify each AXI4 slave independently before full SoC integration.

- Use the exact register map and frozen numerical rules.

- Use the exact FIR coefficient tables from the source-controlled project data.

- Do not add DMA, another master, cache, MMU, interrupt controller, OLED/BLE protocols, or physical hardware dependencies.

- Run unit, AXI4 protocol, integration, end-to-end, and synthesis-oriented checks.

- Do not claim completion until the final acceptance checklist passes.

## 40.1 Final Completion Report

PROJECT_STATUS.md shall record the generated directory tree, RTL modules, testbench/model files, register files, package/library files, firmware, scripts, root run command, tool versions where available, test results, synthesis results, warnings, known limitations, and compliance status.

# 41. Final Kiro-Ready AXI4 Implementation Contract

## 41.1 Implementation Order

| Stage | Required work | Exit condition |
| --- | --- | --- |
| 1 | Read VeeR reference and AXI4 CardioEdge specification | Exact VeeR AXI4 ports/widths/timing documented; no guessed ports |
| 2 | Create cardioedge_pkg and register definitions | Constants/register offsets compile |
| 3 | Create AXI4 memories and interconnect | Address decode, arbitration, AXI handshake and response tests pass |
| 4 | Create AXI4 Timer/UART/GPIO/SPI2 | Unit tests pass |
| 5 | Create AXI4 ADC controller + ADC FIFO + MCP3201 model | Serial waveform and FIFO tests pass |
| 6 | Create FIR input FIFO + FIR + frozen coefficients | Bit-accurate FIR tests pass |
| 7 | Create FIR-QRS FIFO | All flow-control tests pass |
| 8 | Create QRS accelerator | Bit-accurate QRS tests pass |
| 9 | Create interrupt OR | All source combinations pass |
| 10 | Create VeeR AXI4 wrapper/top integration | Processor and peripheral integration compiles |
| 11 | Create firmware/startup/linker/HAL | Firmware builds and initializes system |
| 12 | Create full SoC AXI4 testbench | Boot/configuration/integration tests pass |
| 13 | Run end-to-end regression | Expected ECG results pass |
| 14 | Run synthesis-oriented checks | Synthesizable RTL hierarchy passes |
| 15 | Generate final documentation/status | All deliverables and checklist complete |

## 41.2 VeeR EL2 AXI4 Integration Contract

The VeeR EL2 Programmer's Reference Manual is authoritative for the processor-side AXI4 interfaces. The VeeR core complex provides AXI4 system-bus interfaces for instruction fetch and load/store. The IFU master is read-only and the LSU master supports reads and writes. Both use 64-bit AXI data channels and 32-bit addresses.

VeeR EL2 Core Complex
        |
        +--> IFU AXI4 Master ----+
        |                         |
        +--> LSU AXI4 Master ----+--> veer_el2_axi_wrapper
                                  |
                                  v
                           axi4_interconnect

The wrapper shall expose canonical bus_aw*/bus_w*/bus_b*/bus_ar*/bus_r* signals internally to the interconnect and shall contain all VeeR-specific naming adaptation. The supplied VeeR reference identifies signals such as ifu_axi_awvalid/ifu_axi_arvalid/ifu_axi_rdata and lsu_axi_awvalid/lsu_axi_wdata/lsu_axi_arvalid/lsu_axi_rdata; the exact complete port list shall be copied from the supplied reference during implementation.

## 41.3 Top-Level CardioEdge Interface

| Port | Direction | Width | Purpose |
| --- | --- | --- | --- |
| clk | input | 1 | 50 MHz system clock |
| rst_n | input | 1 | Active-low synchronous reset |
| adc_spi_sclk_o | output | 1 | ADC serial clock |
| adc_spi_mosi_o | output | 1 | ADC serial data output |
| adc_spi_miso_i | input | 1 | ADC serial data input |
| adc_spi_cs_n_o | output | 1 | ADC chip select |
| uart_tx_o | output | 1 | UART transmit |
| uart_rx_i | input | 1 | UART receive |
| gpio_led_o | output | 1 | LED output |
| gpio_buzzer_o | output | 1 | Buzzer output |
| gpio_button_i | input | 1 | Button input |
| spi2_sclk_o | output | 1 | SPI2 serial clock |
| spi2_mosi_o | output | 1 | SPI2 serial output |
| spi2_miso_i | input | 1 | SPI2 serial input |
| spi2_cs_n_o | output | 1 | SPI2 chip select |

## 41.4 FSM Contract

| Block | Required states |
| --- | --- |
| ADC SPI | ADC_IDLE, ADC_START, ADC_CLOCK_LOW, ADC_CLOCK_HIGH, ADC_COMPLETE |
| FIR | FIR_IDLE, FIR_LOAD, FIR_MAC, FIR_SCALE, FIR_SATURATE, FIR_WRITE |
| QRS | QRS_IDLE, QRS_TRACK, QRS_REFRACTORY, QRS_WAIT_BELOW |
| SPI2 | SPI2_IDLE, SPI2_ASSERT, SPI2_TRANSFER, SPI2_COMPLETE |

FSM encoding may be binary or one-hot, but state behavior and transitions shall remain equivalent to the source specification.

## 41.5 Firmware Startup Contract

- Initialize stack/runtime.

- Disable or mask CardioEdge interrupts.

- Clear GPIO outputs.

- Configure Timer LOAD=199999, AUTO_RELOAD=1, IRQ_EN.

- Configure ADC ENABLE=1 and fixed 1 MHz serial clock.

- Drain/clear stale ADC FIFO state.

- Load frozen FIR coefficients and select BASE.

- Configure FIR.

- Configure QRS THRESHOLD=25 and clear QRS statistics.

- Enable QRS interrupts.

- Enable ADC interrupts.

- Enable Timer acquisition.

- Enable global processor interrupt handling.

- Enter the main loop.

## 41.6 Verification Vectors

| ID | Stimulus | Expected result |
| --- | --- | --- |
| V01 | Reset asserted/deasserted | Specified reset values; FIFOs empty; interrupts inactive |
| V02 | Read every mapped register | Correct reset/default value |
| V03 | Write/read writable registers | Value retained subject to side effects |
| V04 | Unmapped AXI read/write | Read zero / write no side effect; transaction completes |
| V05 | Timer LOAD=199999 | Expiry every 200000 cycles |
| V06 | ADC injected sample 0xA55 | FIFO_DATA returns 0xA55 |
| V07 | Fill ADC FIFO to 32 | FULL asserted; 33rd completed sample causes OVERRUN/no overwrite |
| V08 | FIR impulse/reference vector | Output matches frozen reference |
| V09 | FIR max positive/negative stress | Output saturates at 32767/-32768 |
| V10 | FIR input FIFO full | Additional write rejected |
| V11 | FIR-QRS FIFO full | FIR output held; no overwrite |
| V12 | FIR-QRS simultaneous read/write | Count/order remain correct |
| V13 | Known synthetic beat spacing | RR_INTERVAL equals exact sample spacing |
| V14 | RR=250 samples | HEART_RATE=60 |
| V15 | RR=125 samples | HEART_RATE=120 |
| V16 | RR=300 samples | HEART_RATE=50 |
| V17 | Four stable RR values | RR average equals integer average |
| V18 | RR deviation <=25 | No arrhythmia |
| V19 | RR deviation >25 | ARRHYTHMIA_FLAG asserted |
| V20 | Two peaks <50 samples apart | Only first beat accepted |
| V21 | Beat and timer simultaneously | Both interrupt causes observable |
| V22 | ADC sample-ready and overrun simultaneously | Both causes observable/clearable |
| V23 | SPI2 TX while idle | 8-bit transfer completes |
| V24 | SPI2 TX while busy | Write ignored/no corruption |
| V25 | Processor boot | Firmware reaches configured main loop |
| V26 | End-to-end ECG stimulus | Expected beat/RR/BPM/arrhythmia results |
| V27 | Concurrent IFU and LSU AXI requests | Correct arbitration and response IDs |
| V28 | AXI backpressure on each channel | No lost/duplicated transaction |
| V29 | AXI partial WSTRB register write | Only selected byte lanes change |
| V30 | AXI read/write response ordering | Responses match issued IDs and AXI ordering rules |

## 41.7 File-by-File Responsibility

| File/module | Sole responsibility |
| --- | --- |
| cardioedge_top.sv | Top-level integration and external ports |
| veer_el2_axi_wrapper.sv | Only VeeR-to-CardioEdge AXI4 interface adaptation |
| cardioedge_pkg.sv | Common constants, types, addresses, AXI parameters |
| axi4_interconnect.sv | AXI4 routing, arbitration, decode, channel/response routing |
| instruction_memory_axi.sv | Instruction storage and AXI4 read path |
| data_memory_axi.sv | Data storage and AXI4 read/write path |
| timer_axi.sv | 250 Hz timing and timer event |
| uart_axi.sv | UART TX/RX and AXI4 registers |
| gpio_axi.sv | GPIO registers and external GPIO |
| adc_interface_controller_axi.sv | MCP3201 digital SPI controller and AXI4 registers |
| adc_fifo_32x12.sv | 32 x 12 ADC buffering |
| fir_input_fifo.sv | 8 x 16 FIR input buffering |
| fir_filter_31tap.sv | 31-tap fixed-point filtering |
| fir_qrs_fifo_32x16.sv | 32 x 16 FIR-to-QRS streaming |
| qrs_heartbeat_accelerator.sv | Derivative, integration, threshold, RR, BPM, arrhythmia |
| spi2_controller_axi.sv | Generic 8-bit SPI2 master and AXI4 registers |
| interrupt_or.sv | Flat interrupt OR only |
| mcp3201_model.sv | Simulation-only ADC model |
| spi2_sink_model.sv | Simulation-only SPI2 endpoint |
| ecg_stimulus_model.sv | Simulation ECG stimulus |

## 41.8 Final Acceptance Checklist

- Every specified block has an RTL source file.

- Every required AXI4 testbench exists and runs.

- Every required behavioral model exists.

- Every required register file exists and agrees with RTL.

- Every required package/library file exists.

- Every required firmware file exists and builds.

- Every required script exists and is executable where appropriate.

- Root run performs the complete default regression.

- SPECIFICATION.md agrees with this DOCX specification.

- AXI4 signal names are used at all CardioEdge bus boundaries.

- VeeR-specific names are confined to the wrapper.

- All register offsets, fields, access types, reset values, and side effects match.

- All control/data paths match the specification.

- All FSM behavioral states and transitions match.

- All AXI4 error/recovery paths are implemented.

- All interrupts are observable and clearable as specified.

- FIR matches the frozen numerical reference.

- QRS matches the frozen integer reference.

- Firmware sequence matches the specification.

- End-to-end ECG regression passes.

- Synthesis-oriented checks pass.

- No physical hardware is needed.

## 41.9 Final Rule

This document is a deterministic engineering contract. Internal RTL structure may be optimized only when behaviorally equivalent. Any change affecting the AXI4 topology, module boundary, canonical signal, register, memory address, timing contract, FIFO behavior, arithmetic result, interrupt behavior, firmware-visible behavior, or required deliverable requires an explicit specification revision.

# Appendix A — AXI4 Revision Notes

- The supplied source specification was APB-oriented. This document converts the memory-mapped architecture to AXI4 while retaining the original functional requirements.

- The VeeR EL2 reference states that the core complex has separate IFU and LSU system-bus interfaces and supports 64-bit AXI4. The IFU interface is read-only; the LSU interface supports read/write transactions.

- The source memory map and peripheral register map are intentionally unchanged.

- The dedicated FIR-to-QRS stream remains outside the system bus.

- The original APB completion semantics are represented by AXI4 transaction completion; unmapped transactions complete without side effects.

- AXI4-specific implementation details that were not defined by the original CardioEdge document must be verified against the VeeR reference and this AXI4 contract before RTL freeze.

# Appendix B — Source Grounding

This AXI4 specification was prepared from the uploaded CardioEdge specification and the uploaded VeeR EL2 Programmer's Reference Manual. The CardioEdge source defines the functional blocks, address map, registers, FIFOs, numerical processing, interrupts, firmware behavior, verification vectors, and project deliverables. The VeeR reference defines the processor's AXI4 system-bus architecture and native IFU/LSU interfaces.

END OF CARDIOEDGE AXI4 SYSTEM SPECIFICATION v2.0